CONTENTS:

The dialysis process	3

Physical principles of dialysis	3

Diffusion	4

Ultrafiltration	6

Indications for Dialysis -In Chronic Renal Failure	11

A) Absolute indications	11

B) Relative indications	11

C) Early indications	12

D) Specific indications for peritoneal dialysis	12

Indications for Dialysis other than chronic renal failure	13

Dialyzable drugs and Poisons (partial list)	13

Factors to be considered in determine drug’s dialyzability:	14

Factors Determinants for Dialysis	14

Medical Factors	14

Non Medical Factors	14

Contraindications of Dialysis therapy	15

Relative contraindications to dialysis therapy	15

Contraindication for Peritoneal dialysis	15

Absolute	15

Relative Major	15

Relative Minor	15

A. Catheters	16

Veins used:	17

Indication	17

Contraindication of Subclavian and Internal jugular vein catheterization	17

Nursing management (Guide)	17

Bases for cannula selection	18

Complications	18

Management of technical problems	19

Permanent Vascular Access For Hemodialysis	20

Types of permanent V. Access	22

B. Arteriovenous Graft (1974 / 1977)	25

Access Maturation	26

Using Permanent Vascular Access	27

Kind of the needle	27

Placing The Needle in the Access:	27

Removing the needle	28

Care required between dialysis	28

Types of semipermeable  membranes  used for hemodialysis:	29

1-Organic Cellulose (C6H10O5) membranes and its derivatives :	29

2-Synthetic membranes:	30

Dialyzers	31

Coil dialyzer :( are of historic interest)	31

Hollow  Fiber dialyzer	32

Dialyzer specifications	32

	Acetate Dialysate	35

	Bicarbonate Dialysate	35

1. Blood pump	37

2.Dialysate delivery system :	38

5. System disinfection:	43

6. Alarm during Hemodialysis- Problem Solving Guide Lines]	43

1. General Reassessment	48

2.Rinsing and priming	48

4.Obtaining vascular access	48

5.Initiatting dialysis	49

8. Termination of dialysis	50

10. Equipment care	50

b. Mechanical Complications	60

	Management of Medical and Mechanical Complications During Hemodialysis	61

	Disequilibrium syndrome	67

	Intradialytic Hypotension	68

	Membrane Reactions	71

	Muscle cramps during hemodialysis	73

	Skin disorders in dialysis patient	74

	Pruritus (Itching)	75

.	77

	Causes of Chest pain in ESRD  patients	77

.	78

	Hemodialysis associated angina	78

	Air Embolism	79

	Ruptured  Dialyzer	81

	Clotted dialyzer	81

	Needle-Site Bleeding	82

	High Venous Pressure	83

	Abnormal Arterial Pressure	84

	Hemolysis	85

	Conductivity	86

Principles of dialysis

# The dialysis process

The concept of dialysis, that is partition (separation) of substance between two solutions by use of semipermeable membrane, is quite simple. Extracorporeal dialysis employs the artificial kidney (dialyzer) as semipermeable membrane, while   Intracorporeal   dialysis employs the peritoneal membrane.

# Physical principles of dialysis

Osmosis are the basic physical principles of dialysis. Diffusion is the net directional movement of molecules occurring from a solution of higher concentration to a solution of low concentration. Ultrafiltration is the movement of solvent across a semipermeable membrane in response to a pressure difference applied across the membrane.If the solutes dissolved in the solvent is small enough to permeate the membrane, they are dragged along with the solvent and cross over to the other side, and this called Convection. Osmosis the movement of the solvent (e.g. water) from the side of low concentration to the side of higher concentration.

# Diffusion

If two solution of different composition are placed on different side of a semipermeable membrane, solutes will move from the solution with the highest concentration to the solution with the lowest concentration.The rate of movement will depend on:-

- The concentration gradient for the solute between the two solutions

- Permeability of the membrane to the solute

- Surface area of the membrane.

- The size of the solute is highly correlated with its molecular weight. The heavier, larger solute moves more slowly along the concentration gradient than smaller lighter solutes. Thus, dialysis is most effective in removing small solutes and less effective in removing larger solutes, particularly those over 1000 Dalton.

affects the diffusive transport of solute across the membrane by two different mechanisms:

A. the binding of the solute to protein: Diffusible substance may bind to proteins forming dialysis membrane impermeable complex. Such solutes are no longer available for diffusion (40-50% of measured calcium in patient blood is available for diffusion The percentage of the total concentration of a diffusible solute (actually free to diffuse) is described as “solute activity”.

Molecular weight in Dalton of some nonionic substances

| Substance | Molecular weight in Dalton |
| --- | --- |
| BUN Urea Creatinine Ethanol Methanol Glucose Vit B12 Albumin | 28 60 113 46 33 180 1355 68000 |

B. the Gibbs-Donnan effect: Blood proteins are dialysis membrane impermeable, negatively charged and tend to accumulate at the membrane surface during dialysis. Coresponding numbers of membrane permeable cations such as sodium, calcium, magnesium must then retain in the blood to preserve elecroneutrality. This results in imbalance in the concentration of ions across the dialysis membrane. The protein-induced ion transport asymmetry is called the “Gibbs-Donnan effect”. It indirectly affects the magnitude of the concentration gradients required to drive diffusion across dialysis membrane.

As solute movement continues over a peroid of time, the concentration falls in the solution of higher concentration, rises in the solution with the lower concentration, and the two solutions approach each other in composition i.e Equilibrium. As a result of this dissipation of the concentration gradient, the transfer of solute slows with the time.The maximum rate of solute transfer occurs initially when the concentration gradient is greatest.

No hydrostatic pressure is applied

[Van Stone et al 1994, Physiology of peritoneal dialysis in handbook of dialysis 2nd edition]

The dissipation of the concentration gradient can be minimized and the transfer of solute optimized by increasing the volume of the fluids.

a) In a static system (comparable to peritoneal dialysis) this is accomplished by replacing the recipient fluid with fresh solution at periodic intervals.

b) In a flowing system (comparable to hemodialysis) this accomplished by increasing the flow rate of parent fluid (blood) or recipient fluid (dialysate).

Both artificial and natural membranes are more permeable to small solutes than large solutes. Thus, dialysis is most effective in removing small solutes and less effective in removing larger solutes, particularly those over 1000 daltons.The surface area of the membrane available for diffusion affects the amount of solute transferred.

Mechnisms of solute removal in intermittent hemodialysis and continuous renal replacement therapy

|  | IHD | CRRT |
| --- | --- | --- |
| Small solutes (lees than 300 D) | Diffusion | Convection (CVVH) Diffusion (CVVHD) |
| Middle molecules (500-5000) | Diffusion Convection | Convection Diffusion |
| LMW protein (5000-50 000) | Convection Diffusion Adsorption | Convection Adsorption |
| Large molecules (m0re than 50 000 | Convection | Convection |

# Ultrafiltration

The process of water removal from the blood stream is called ultrafiltration; the fluid removed is the ultrafitrate.  The UF during dialysis is performed for the purpose of removing water accumulated by ingestion of fluid or by metabolism of food during the interdialytic period. It is essential to prescribe and control the fluid removal rate so that total fluid removed during dialysis will be equal to the total fluid gained since the previous dialysis or from the dry weight.

Schematic diagram of osmotic ultrafiltration

No hydrostatic pressure is applied. Triangles represent osmotic agent introduced to right compartment. This causes an early water shift to the right (ultrafitration), but this is later reversed if the osmotic agent is also small enough to diffuse along the concentration gradient from right to left. Thus only solutes of such size that they do not easily permeate the semipermeable membrane are capable of exerting a sustained osmotic force.   [Sorkin MI et al 1994, Physiology of peritoneal dialysis in handbook of dialysis 2nd edition]

Schematic diagram of hydrostatic  ultrafiltration

The application of external pressure forces movement of water from left to right.  Low molecular weight constituents will be swept through the membrane with this water (solvent drag.). In dialysis setting the pressure gradient is generated by manipulation if dialysis fluid parameters such as pressure volume or flow.

#### Mechanism of Ultrafiltration

A) In Hemodialysis:

1-Hydrostatic pressure

The primary driving force for ultrafiltration is the hydrostatic pressure difference across the membrane, which is the Transmembrane Pressure (TMP), expressed in millimeters mercury (mmHg). The TMP is determined by the average or mean pressure in the blood compartment minus the mean dialysate compartment pressure.The relationship of ultrafiltrate to TMP is entirely dependent on the membrane (Dialyzer) properties. The permeability of dialyzer membranes to water is high, varies consonsiderably, and is a function of membrane thickness and pore size.The total capacity of the dialyzer for ultrafiltration is given by the Ultrafiltration Coefficient (KUF).

The KUF is defined as the number of milliliters of fluid per hour that will be transferred across the membrane per mmHg pressure gradient across the membrane The KUF of most dialyzer ranges from 2 to 6 ml/hour. The relationship between ultrafiltration, KUF and TMP is expressed as:

Ultrafiltration rate (ml/hr) = KUF X TMP

How do you calculate UF rate from KUF?

If one needs to remove 2 kg during a period of 4 hours

1- Add the volume of saline that will be given at the end of dialysis (usually 300 ml) and the amount of ingested fluid during dialysis (e.g., 100ml).

2- This means that 2.4 L will have to be removed during 4 hours dialysis i.e.2.4 X 1000 /4 = 600 ml/hour.

3- When using a dialyzer with KUF value of 4.0 ml/hour, the TMP will need to be set at 600/4 =150 mm Hg. N.B. (If the dialyzer KUF

2-Osmotic Ultrafiltration:

Osmotic ultrafiltration does play an indirect role in total ultrafiltration; water shifts from intracellular to the extracellular compartment which occur during hemodialysis (so-called plasma refilling) can be optimized by introduction of an effective concentration of an osmotic agent into the extracellular space. Sodium is employed in some dialysis practice especially during sodium profiling

Indications and contraindications of dialysis

# Indications for Dialysis -In Chronic Renal Failure

In patients with chronic renal failure factors to be considered before initiating dialysis should include comorbid conditions and patient preference. Timing of therapy is dictated by serum chemistries and symptoms

## A) Absolute indications

Uremic Pericarditis

Uremic Encephalopathy or Neuropathy

Pulmonary edema (unresponsive to diuretics)

Severe Hypertension

Severe hyperkalemia

Intractable acidosis

Severe Bleeding diathesis

Persistent gastrointestinal symptoms

S.Creatinine more than 12 mg/dl, BUN more than 100 mg/dl

## B) Relative indications

Mild encephalopathy or neuropathy

Severe edema (unresponsive to diuretics)

Progressive gastrointestinal symptoms

Recurrent GI “itis”: stomatitis, gastritis, dudenitis, pancreatitis

Ascitis without hepatic disease

Anemia refractor to Erythropoietin

Mild Bleeding diathesis

Pruritus

Infectious complications

Depression

## C) Early indications

Decrease ideal body weight

Decrease in muscle mass (decrease s creatinine or its clearance)

Decrease in s.albumin to less than 4 g/l

GFR less than 15 ml/min (by I iothalamate)

S. Creatinine >10 mg/dl and bun >100 mg /dl

Decrease in s.transferrin

Low total cholesterol

Growth retardation in children

## D) Specific indications for peritoneal dialysis

Patients with cardiovascular or hemodynamic instability

Hemodialysis patients with vascular access failure or can not be created (e.g. diabetic patients)

High risk of anti coagulation

Patients in the older age group (over 65) and small children

Severe hemodialysis-related symptoms or disequilibrium

Social reason

# Indications for Dialysis other than chronic renal failure

1-Acute renal failure

2-Poisons and Drug intoxication

3-Hypercalcaemia

4-Hyperuricemia

5-Hypothermia

6-Metabolic alkalosis (special dialysis solution required)

## Dialyzable drugs and Poisons (partial list)

a) Barbiturates            : Phenobarbital -Pentobarbital -Amobarbital

b) Alcohol's                 : Methanol -Ethanol -Ethylene glycol -Isopropanol

c) Analgesics               : Acetylsalicylic acid Methylsalicylate

d) Metals                     : Calcium -Potassium -Sodium -Lithium

e) Endogenous toxins  : Uric acid -Uremic toxins -Hyperosmolar state

f) Halides                     : Bromide

e) Miscellaneous          : Theophylline -Mannitol -Radiocontrast -Thiocynate -

Boric acid -Aniline

NB   1) Dialysis for poisoning should be considered only when supportive measures are                                                                   ineffective or there is impending irreversible organ toxicity.

2) Hemoperfusion is required in some cases

### Factors to be considered in determine drug’s dialyzability:

#### 1) Dialysis related factors:

Dialyzer membrane characteristics

Surface area

Blood flow rate

Dialysate flow rate

Degree of ultrafitration

Duration of dialysis

#### 2) Drug related factors   :

Availability of the drug in the plasma

Drug pharmacokinetic characteristics

Drug molecular weight (<500 cross the membrane more readily)

Drug solubility (water soluble are dialyzable than lipid soluble

# Factors Determinants for Dialysis

Factors determining the mode of chronic dialysis should include medical and non medical factors which have an impact on the treatment modality. Physicians have the responsibility to discuss the therapeutic options and offer their advice and recommendations about the choices. In general renal transplantation should be recommended as the preferred mode of renal replacement therapy in whom surgery and immunosupression is safe and feasible.

## Medical Factors

- Age

- Comorbid medical illnesses

- Patient survival

- Patient rehabilitation

- Quality of life

## Non Medical Factors

- Government-imposed Economic limitations

- Physician  and Patient bias

- Resource availability

- Social ,Religious ,Cultural mores

- Availability of transplantation

- Family support

- Cost , race , sex , reimbursement

# Contraindications of Dialysis therapy

Principally there is no absolute contraindication to dialysis therapy. Advanced age in and of itself is not a contraindication to dialysis therapy. Many elderly are physiologically equivalent to young patients.

## Relative contraindications to dialysis therapy

- Advanced malignancy (except multiple myeloma)

- Alzheimer’s disease

- Multi-infarct dementia

- Hepatorenal syndrome

- Advanced liver cirrhosis with encephalopathy

- Hypotension unresponsive to pressors

- Terminal illness

- Organic brain syndrome

## Contraindication for Peritoneal dialysis

### Absolute

- Peritoneal fibrosis

- Pleuroperitoneal leak

### Relative Major

- Chronic Ostomies

- Severe hypercatabolic state

- Fresh aortic prosthesis

- Recent Abdominal surgery

- Recent Thoracic surgery

- Extensive Abdominal adhesions

- Quadriplegia

- Blindness

- Physical handicaps

- Mental Retardation

### Relative Minor

- Polycystic Kidney disease

- Diverticulosis

- Obesity

- Peripheral vascular disease

- Hyperlipidemia

- Social

Vascular access

Acute Hemodialysis Vascular Access–Noncuffed Catheters (Guide 2 DOQI)

A. Hemodialysis access of less than 3 weeks’ duration should be obtained using a noncuffed, or a cuffed, double-lumen percutaneously inserted catheter.

B. These catheters are suitable for immediate use and should not be inserted before needed.

C. Noncuffed catheters can be inserted at the bedside in the femoral, internal jugular, or subclavian position.

D. The subclavian insertion site should not be used in a patient who may need permanent vascular access.)

E. Chest x-ray is mandatory after subclavian and internal jugular insertion prior to catheter use to confirm catheter tip position at the caval atrial junction or the superior vena cava and to exclude complications prior to starting hemodialysis.

F. Where available, ultrasound should be used to direct insertion of these catheters into the internal jugular position to minimize insertion-related complications.

G. Femoral catheters should be at least 19-cm long to minimize recirculation. Noncuffed femoral catheters should not be left in place longer than 5 days and should be left in place only in bed-bound patients.

H. Nonfunctional noncuffed catheters can be exchanged over a guidewire or treated with Urokinase as long as the exit site and tunnel are not infected.

I. Exit site, tunnel tract, or systemic infections should prompt the removal of noncuffed catheters

Temporary vascular access:

A. Catheters             B.Quinton-Scribner external arteriovenous shunt

## A. Catheters

1. Single lumen, 1961

2. Double lumen, 1980s

These catheters made of Teflon, Silastic, or Polyurethane

### Veins used:

Femoral vein                 Subclavian vein                     Internal jugular vein

### Indication

1. Acute renal failure

2. ESRF patients whose permanent access is immature

3. Patients on peritoneal dialysis requiring temporary hemodialysis

4. Patient for plasmapheresis

5. Patients for venovenous or arteriovenous continuous renal replacement therapy.

### Contraindication of Subclavian and Internal jugular vein catheterization

1. Patients with acute respiratory distress (cannot be supine or in Trendelenberg position. 2. Patients with subclavian vein stenosis

3. Abnormal coagulation parameters

### Nursing management (Guide)

        All of these catheters are prone to infection therefore aseptic technique when initiating and terminating dialysis is of outmost importance.

        The caps and ports should be wrapped in a dressing soaked in Betadine for 5 minutes before initiating and ending dialysis.

        The exit site should be cleaned with Peroxide and Betadine after every treatment, and a sterile dressing should be applied. The appearance of the exit site especially redness or drainage must be documented

        Careful heparinization is necessary and should be guided by manufacturer's instruction and according to the size (or dead space) of each catheter. Indwelling heparin should be aspirated before dialysis to avoid excessive heparinization. A preferable method of maintaining catheter patency is the instillation of 1000-5000 U of heparin into the catheter immediately after dialysis. The volume necessary to deliver heparin depends on the size (dead space) of the catheter. Clamp the catheter firmly and once to avoid blood suction into the tip of the catheter by dropper action

        Do not attempt to instill saline into a clotted catheter. This will force the clot into the vascular system.

        Minimal handling of the catheter is strongly advocated, and then only experienced personnel. If the catheter is needed for parentral nutrition or blood transfusion a Leuer-lok connections should be used to prevent air embolus or blood loss.

### Bases for cannula selection

1. Subclavian and internal jugular vein catheter can be lift in place for several weeks while femoral catheter almost always removed within 3 days.

2. Once subclavian and internal jugular vein catheters patient can be dialyzed as an outpatient while patient with femoral catheter must be hospitalized.

3. Femoral catheter is easily insertable while subclavian and internal jugular vein catheters insertion requires a skilled operator.

4. Complication from femoral catheterization usually minor and never life threatening in contrary to subclavian and internal jugular vein catheters insertion.

5. New femoral catheter insertion for each dialysis is the best approach in bacteraemic patients

### Complications

|  | Femoral | Subclavian | Jugular |
| --- | --- | --- | --- |
| Early | Pain at insertion site-Bleeding-Puncture of the femoral artery, Intestinal perforation | Pain at insertion site-Bleeding- Puncture of subclavian artery-Pneumothorax-Hemothorax- Brachial plexus injury--Puncture of superior vena cava-Mediastinal haemorrge Peric.temponade-Arrhythmias | Pain at insertion site-Bleeding- Puncture of carotid artery-Pneumothorax-Hemothorax- Arrthmias |
| Late | Groin hematoma-Retroperitoneal hematoma-Infection-Catheter clotting | Infection-Catheter clotting-Subclavian vein thrombosis or stricture | Infection -Catheter clotting Thrombosis and stricture are quite low incidence |

### Management of technical problems

1. Poor blood flow-Lowering patient's head if using subclavian or Jugular vein

-Turning patient's head to opposite side if using subclavian or Jugular V

-Apply external pressure to the exit site

-Rotate the catheter shaft 180 degrees

-Reverse the lines (last resort)

-Replace the catheter if the catheter has been in place for less than

week2. Clotting            -Proper heparinization postdialysis

-Clamp the catheter firmly and once to avoid blood suction into the tip

of the catheter by dropper action

-Finrinolytic agents such as urokinase or streptokinase

-Replace the catheter if the catheter has been in place for less than

week.

3. Kinking            -Replace the catheter if the catheter has been in place for less than

week.

.

Tunneled Cuffed Catheter Placement: Type and Location: (DOQI)

A. Tunneled cuffed venous catheters are the method of choice for temporary access of longer than 3 weeks’ duration. (They also are acceptable for access of shorter duration.) In addition, some patients who have exhausted all other access options require permanent access via tunneled cuffed catheters. For patients who have a primary AV fistula maturing, but need immediate hemodialysis, tunneled cuffed catheters are the access of choice. Catheters capable of rapid flow rates are preferred. (Evidence/Opinion)

B. The preferred insertion site for tunneled cuffed venous dialysis catheters is the right internal jugular vein. Other options include: the right external jugular vein, the left internal and external jugular veins, subclavian veins, femoral veins, or translumbar access to the inferior vena cava. Subclavian access should be used only when jugular options are not available. Tunneled cuffed catheters should not be placed on the same side as a maturing AV access, if possible. (Evidence)

C. Fluoroscopy is mandatory for insertion of all cuffed dialysis catheters. The catheter tip should be adjusted to the level of the caval atrial junction or into the right atrium to ensure optimal blood flow. (Atrial positioning is only recommended for catheters composed of soft compliant material, such as silicone.) (Opinion)

D. Real-time ultrasound-guided insertion is recommended to reduce insertion-related complications. (Evidence/Opinion)

E. There is currently no proven advantage of one cuffed catheter design over another. Catheters capable of a rapid blood flow rate are preferred. Catheter choice should be based on local experience, goals for use, and cost. (Evidence/Opinion)

# Permanent Vascular Access For Hemodialysis

| Patient Evaluation Prior to Access Placement |
| --- |
| Consideration | Relevance |
| Patient History |  |
| History of previous central venous catheter | Previous placement of a central venous catheter is associated with central venous stenosis. |
| Dominant arm | To minimize negative impact on quality of life, use of the nondominant arm is preferred. |
| History of pacemaker use | There is a correlation between pacemaker use and central venous stenosis. |
| History of severe congestive heart failure | Accesses may alter hemodynamics and cardiac output. |
| History of arterial or venous peripheral catheter | Previous placement of an arterial or venous peripheral catheter may have damaged target vasculature. |
| History of diabetes mellitus | Diabetes mellitus is associated with damage to vasculature necessary for internal accesses. |
| History of anticoagulant therapy or any coagulation disorder | Abnormal coagulation may cause clotting or problems with hemostasis of accesses. |
| Presence of comorbid conditions, such as malignancy or coronary artery disease, that limit patient’s life expectancy | Morbidity associated with placement and maintenance of certain accesses may not justify their use in some patients. |
| History of vascular access | Previously failed vascular accesses will limit available sites for accesses; the cause of a previous failure may influence planned access if the cause is still present. |
| History of heart valve disease or prosthesis | Rate of infection associated with specific access types should be considered. |
| History of previous arm, neck, or chest surgery/trauma | Vascular damage associated with previous surgery or trauma may limit viable access sites. |
| Anticipated renal transplant from living donor | Temporary access may be sufficient. |
| Physical Examination |  |
| Physical Examination of Arterial System |  |
| Character of peripheral pulses, supplemented by hand-held Doppler evaluation when indicated | An adequate arterial system is needed for access; the quality of the arterial system will influence the choice of access site. |
| Results of Allen test | Abnormal arterial flow pattern to the hand may contraindicate the creation of a radial-cephalic fistula. |
| Bilateral upper extremity blood pressures | Pressures determine suitability of arterial access in upper extremities. |
| Physical Examination of Venous System |  |
| Evaluation for edema . | Edema indicates venous outflow problems that may limit usefulness of the associated potential access site or extremity for access placement |
| Assessment of arm size comparability | Differential arm size may indicate inadequate veins or venous obstruction which should influence choice of access site. |
| Examination for collateral veins | Collateral veins are indicative of venous obstruction. |
| Tourniquet venous palpation with vein mapping | Palpation and mapping allow selection of ideal veins for access. |
| Examination for evidence of previous central or peripheral venous catheterization | Use of central venous catheters is associated with central venous stenosis; previous placement of venous catheters may have damaged target vasculature necessary for access. |
| Examination for evidence of arm, chest, or neck surgery/trauma | Vascular damage associated with previous surgery or trauma may limit access sites. |
| Cardiovascular Evaluation |  |
| Examination for evidence of heart failure | Accesses may alter cardiac output. |

Arteriography is useful to avoid extremity ischemia in patients with diminished pulses in whom access in the extremity is still desired. However, the Work Group concluded that arteriography is only rarely required.

Venipucture in the non-dominant forearm should be minimized in patients with chronic renal failure

Permanent V. Access should be created when creatinin clearance drops to ~ 15 ml/minute or 3-6 months prior to initiation of hemodialysis.

The risk of subclavian vein stenosis in patient who have had previous subclavian catheter is high therefore angiographic assessment is recommended prior to access placement.

## Types of permanent V. Access

A.Arteriovenous Fistula                          B.Arteriovenous Graft

A.Arteriovenous Fistula (1966)

Description:

The arteriovenous fistula consists of surgical anastomosis of an adjoining artery and vein. The diversion of the arterial blood causes the used veins to become enlarged and prominent and stronger because of the greater flow of the arterialized blood through them.

Selection of Permanent Vascular Access and Order of Preference for Placement of AV Fistulae (Guide 3 DOQI)

A. The order of preference for placement of AV fistulae in patients with kidney failure who will become hemodialysis dependent is:

1. A wrist (radial-cephalic) primary AV fistula (Evidence)

2. An elbow (brachial-cephalic) primary AV fistula (Evidence/Opinion)

B. If it is not possible to establish either of these types of fistula, access may be established using:

1. An arteriovenous graft of synthetic material (eg, PTFE) (Evidence) or

2. A transposed brachial basilic vein fistula (Evidence) C. Cuffed tunneled central venous catheters should be discouraged as permanent vascular access.

Types and blood vessels used:

Radiocephalic Fistula       : Anastomosis of Radial artery   and Cephalic vein

Brachiocphalic Fistula      : Anastomosis of Brachial artery and Cephalic vein

BrachioBasailic Fistula     : Anastomosis of Brachial artery and Basilic vein

Ulner artery is infrequently used

Anastomosis:

End of the vein to side of the artery

End of the vein to end of the artery

Side of the vein to side of the artery

Surgical Technique: The details of the surgical technique are beyond the scope of the Tips

Postoperative Care:

1. Elevate the arm to minimize edema.

2. Avoid tight dressings

3. Check the fistula daily for, thrill, hematoma, and evidence of ischemia

4. After 4-5 days some exercises could be started

a- A tourniquet may be placed around the upper arm to cause distention of the veins, leave it for about 30 minutes and may be repeated several time each day.

b- Hand exercise, such as squeezing rupper ball, while the tourniquet is in place.

c- Warm compresses to speed the venous distention.

5. Arteriovenous fistula could be used after 2 months.

### B. Arteriovenous Graft (1974 / 1977)

Description: A biologic, semi biologic or prosthetic graft implanted subcutaneously and attached to an artery and vein. It is used in patients who do not have adequate vessels to create an arteriovenous fistula. It can be used after 2-3 weeks.

Type and Location of Dialysis AV Graft Placement

If a primary AV fistula cannot be established, a synthetic AV graft is the next preferred type of vascular access Grafts may be placed in straight, looped, or curved configurations. Designs that provide the most surface area for cannulation are preferred. Location of graft placement is determined by each patient's unique anatomical restrictions, the surgeons's skill, and the anticipated duration of dialysis)

1. Subcutaneous autogenous vein grafts. A segment of patient's own vein is attached to an artery a vein and used for dialysis after becomes prominent and healing.

2. Bovine grafts: Carotid artery from cattle is used after special processing.

3. Synthetic grafts: Many synthetic materials are available [Macron, Polytetfluoroethylene (PTFE)] inserted surgically and require 2 weeks to mature. Polytetrafluoroethylene (PTFE) tubes are preferred over other synthetic materials.

4. PTFE graft with transcutaneous device for needle free access. The system is accessed through self-sealing device with locking ring. A special set attaches to the blood tubing.

Surgical Technique: The details of the surgical technique are beyond the scope of the Tips

Postoperative care: As in Arteriovenous fistula

## Access Maturation

Guide 9 DOQI

A. A primary AV fistula is mature and suitable for use when the vein’s diameter is sufficient to allow successful cannulation, but not sooner than 1 month (and preferably 3 to 4 months after construction.)

B. The following procedures may enhance maturation of AV fistulae:

1. Fistula hand-arm exercise (eg, squeezing a rubber ball with or without a lightly applied tourniquet) will increase blood flow and speed maturation of a new native AV fistula. (Opinion)

2. Selective obliteration of major venous side branches will speed the maturation of a slowly maturing AV fistula. (Opinion)

3. When a new native AV fistula is infiltrated (ie, presence of hematoma with associated induration and edema), it should be rested until swelling is resolved

C. PTFE dialysis AV grafts should not routinely be used until 14 days after placement. Cannulation of a new PTFE dialysis AV graft should not routinely be attempted, even 14 days or longer after placement, until swelling has gone down enough to allow palpation of the course of the graft. Ideally, 3 to 6 weeks should be allowed prior to cannulation of a new graft.

D. Patients with swelling that does not respond to arm elevation or that persists beyond 2 weeks after dialysis AV access placement should receive a venogram or other noncontrast study to evaluate central veins.

E. Cuffed and noncuffed hemodialysis catheters are suitable for immediate use and do not require maturation time. (Evidence)

## Using Permanent Vascular Access

Poor vascular access is a limiting factor to patient survival on hemodilysis. Therefore great care must be taken to maintain adequate vascular access.

### Kind of the needle

Sixteen, fifteen and fourteen gouge needles are used for hemodialysis. Smaller diameter (higher gouge) needles seriously limit the blood flow rate.

Higher flow rate may be possible by using bigger diameter needle, but at considerable increase in negative pressure which increased the possibility of sucking air into the system or damaging the intima of vessels. Resistance to the flow occurs in long needles, therefore the shortest practical needle is desirable.

The selection of the needle depends on:

1-Amount of subcutaneous tissue to be penetrated

2-Size of the vein (access).

3-Angulation of the vein.

### Placing The Needle in the Access:

1-Aseptic technique is essential.

2-Local anesthesia may be used in some patient.

3-Localize the fistula or graft; depth ; angulations ; maximum thrill ; and site of insertion ; hence the angle of needle insertion is decided (~ 45 degree).

3-Insert the inlet (Arterial) needle proximal to the fistula or close to arterial anastomosis of the graft by 3 cm at least, to avoid intimal damage and the subsequent thrombosis.

4-The return ( venous )needle should inserted pointing toward the heart approximately 5 cm proximal to the arterial needle. This opposite direction meant to avoid recirculation. Such event may be undetected , the patient gets poor dialysis.

A-Black blood syndrome : When recirculation is quite severe the blood  becomes very acidic (pH <7 ) the RBCs cannot carry oxygen and the blood appears very

B-If venous needle cannot be inserted ,often vein can be found in another limb or single needle technique to be used.

C-Insert the needles at least 2 cm or more from the previous sites each time to ensure complete healing of the vein.

D-Initiate heparenization after insertion of both needle,.( to avoid hematoma ).

### Removing the needle

It is important to maintain adequate pressure either by hand or tight pressure dressing  over the puncture site for 15-20 minutes after needles is removed.

### Care required between dialysis

1-Good hygienic condition is important. Instruct the patient to wash fistula arm with water and soap predialysis.

2-Advise the patient to remove the dressing  few hours after dialysis .

3-Advise the patient to avoid trauma to the access or to sleep on the same arm.

4-Educate the patient to:

a)Feel the bruit over the fistula (Touch)

b)Observe for signs of infection; redness, pain, swelling, exudates (Look).

What to say to the patient to protect his access?

•Make sure your nurse or technician checks your access before each treatment.

•Keep your access clean at all times.

•Use your access site only for dialysis.

•Be careful not to bump or cut your access.

•Don't let anyone put a blood pressure cuff on your access arm.

•Don't wear jewelry or tight clothes over your access site.

•Don't sleep with your access arm under your head or body.

•Don't lift heavy objects or put pressure on your access arm.

## Types of semipermeable  membranes  used for hemodialysis:

### 1-Organic Cellulose (C6H10O5) membranes and its derivatives :

Cellulose is the most common type of dialyzer membrane. It is  obtained from treated wood products and cotton with heat and chemicals and formed into sheets or extruded through dies as hollow fibers ;e.g. regenerated cellulose, cuprammonium cellulose (cuprophan), cuprammonium rayon, saponified cellulose ester. Cellulose acetate and triacetate are other widely used cellulose substances.

Nature: The membrane is visualized as a thin sheet with tiny pores. These pores are small enough to hold back blood cells and plasma proteins, yet large enough to permit water molecules and many solutes to pass through. Selective permeability is improved by making the membrane thinner, increasing the numbers of channels between fibers , or increasing the diameter of passages.

Positive and Negative features: Hollow fibers have minimal compliance and permit precise TMP calculation. They are inexpensive. Cellulose membranes however have more incompatibility problems than do synthetic membranes specially cuprophan.

### 2-Synthetic membranes:

Synthetic membranes include

Polyacrylonitrile (AN ,PAN)

Polysulfone

Polycarbonte polymide

Polymethylmethacrylate (PMMA)

Ethylene-vinyl alcohol copolymer

Nature: They are a thin , smooth luminal surface, supported by a sponge-like wall structure. All have ultrafiltration coefficients of 20 to 70 ml/hr/mmHg or more.

Positive features: Middle molecule clearance is greater than cellulose membranes. B. Microglobulin is removed by adsorption to the membrane 100 mg /treatment. They have better biocompatibility with blood than cellulose membrane.

Negative features: Very expensive. Automated ultrfiltration control is required because of very high water permeability Significant protein loss by adsorption. Backfiltration from the dialysate and risk of bacterial or endotoxin contamination, due to high hydraulic permeability

Membrane biocompatibility:

Material used in manufacture of dialysis membrane is associated with some degree of blood-material interaction. Complement activation that occurs during dialysis with cellulose membrane is instigated by: 1)Free hydroxyl radicals on the membrane surface. 2) A cellulose material called Limulus-Amebocyte-Lysate-Reactive-Material (LALRM). This activation leads to release of histamine, thromboxane, interleukin 1, and tumor necrosis factor as well as direct action on blood cells. The clinical manifestation varies from minimal symptoms to severe anaphylactic reaction. (See membrane reactions)

# Dialyzers

Types of Dialyzers (Artificial kidneys)

Dialyzers consists of a series of parallel flow paths designed to provide a large surface contact area between blood and dialysate. The dialyzing membrane has pores varying between 11-30 um. The membranes are protected  by dialyzer shell with 4 ports ,2 for blood and 2 for dialysate  Dialyzers are broadly classified as:

1-Coil dialyzer       2-Parallel Plate dialyzer       3-Hollow Fiber dialyzer

### Coil dialyzer :( are of historic interest)

Basically consists of a flattened cellulose tubing  wrapped as a coil and through which patient’s blood flow during dialysis.The blood channels was long to obtain the needed surface area, and resistance was high. UF was unpredictable and blood leak were frequent.

Parallel Plate dialyzer

Structure:

-Sheets of membranes are placed between supporting plates The plates have ridges and grooves to support the membrane and allow flow of dialysate along it.

-Resistance to blood flow is low.  The surface area vary from 0.25 to 1.5 msq.

Advantages:

-Blood volume is about 50-100 ml at 100 mmHg increases with high TMP (bulging sheets)

-Heparin requirement usually low, minimal clotting in the blood compartment.

-Ulttrfiltration is reasonably predictable and controllable.

Disadvantages:

-Formation of local thrombi around inlet and outlet ports and corners due to uneven blood flow at these parts. These may lead to bacterial growth and endotoxin formation , therefore plates are not often reused.

### Hollow  Fiber dialyzer

Structure:

It consists of numerous hollow fibers (capillaries) through which the patient’s blood flow. The hollow fibers are tiny its diameter ~150-250 um. The wall thickness as little as 7 um. The number of fibers may be 20,000 or more ,depending upon length ,kind of membrane , and surface area of the dialyzer. Hollow fiber is the most commonly used dialyzers

Advantages:

-It contained low blood volume in relation to surface area 60-90 ml (low priming volume).

-Resistance to blood flow is low.

-Ultrafiltration can be precisely controlled,

-They are well adapted to reuse.

Disadvantages:

-Deaeration of the fiber predialysis is necessary to prevent air lock of the fibers.

-Uneven blood distribution at the inflow header space leads to relative stagnation centrally, with reduced perfusion and clotting of some center fibers with the subsequent high residual blood volume and aggravation of anemia.

-More heparin is required for most of the patients.

-Adverse patient reaction due to residual toxic products of sterilizing agent.

## Dialyzer specifications

Data about the dialyzer includes the following  information which guides the dialysis personnel to select proper dialyzer for each patient :

1-Type of the dialyzer :(see the previous tip)

2-Material of the membrane :(see the previous tip)

3-Ultrafiltration Coefficient ,(KUF). (see the previous tip for Ultrafitration)

The KUF is defined as the number of milliliters of fluid per hour that will be transferred across the membrane per mmHg pressure gradient across the membrane The KUF of most dialyzer ranges from 2 to 6 ml/hour. If KUF is low (or high) the permeability to water is low (or high). In vivo KUF is often lower by 5-30% than in vitro value The relationship between ultrafiltration, KUF and TMP is expressed as:

Ultrafiltration rate (ml/hr) = KUF X TMP

4-Clearance  Usually reported at blood flow rates of 200,300,and 400 ml/minute

Urea :  A high efficiency dialyzer with thin, large surface area, wide pores, good design will remove a higher percentage of waste products. The efficiency of a dialyzer in removing urea can be described by a constant referred to as Mass transfer urea coefficient -KoA .This constant influence the relation between the blood flow rate to the clearance. clearance .Dialyzers of low efficiency have in vitro KoA value of less than 300 ; used for acute dialysis  and small  patient. Dialyzers of moderate efficiency have in vitro KoA value of 300-600;higher efficiency dialyzers have KoA more than 600-700.Once KoA of the dialyzer is known a nomogram can be used to predict the blood water urea clearance (Kw) at given blood (Qb) and dialysate (Qd) flow rate. The in vitro KoA is usually overestimated ,therefore  in vivo value should be estimated from another nomogram.

Creatinine ,Vit.B12: The dialyzer creatinine clearance is ~ 80% of urea clearance. Vitamin B12 dialyzer clearance range is 30-60 ml/minute. It is used as indication for clearance of higher molecular weight.

Molecular weight of some nonionic substances in mol wt:

BUN     = 28      Urea      = 60        Creatinine = 113         Albumin  = 68000

Glucose =180     Vit B12 = 1355    Ethanol     = 46           Methanol = 33

5-Surface area of the dialyzer:

Normally large surface area of the dialyze have high urea clearance. However dialyzer design and the thinness of the membrane are quit important. Large surface area is an undesirable feature when an unsubstituted cellulose membrane is used. It increases the degree of complement activation.

6-The priming volume :

It is related to the membrane surface area; usually 60-90 ml in hollow fibers. If this added to the priming volume of the blood lines(100-150 ml), the total extracorporeal circuit will be 160-270 ml

7-Mode of dialyzer Sterilization

1-Exposure of the dialyzer to Ethylene Oxide gas

2-Gamma irradiation - becoming popular

3-Steam autoclaving - becoming popular

Dialysate

The term dialysate refers to the fluid and solute that have crossed a membrane i.e. effluent dialysis fluid. However, in current renal jargon all dialysis fluid, fresh or used is called dialysate. Dialysate is sometimes called “bath". The composition of dialysate is nearly correspond  to that of plasma water.

- Composition and function of dialysate

Composition of dialysate:

There are five compounds

1. Sodium chloride

2.Sodium acetate or Sodium bicarbonate

3. Calcium chloride

4. Potassium chloride

5. Magnesium chloride.

Glucose may be included in some formulas.

The use of bicarbonate in concentrate causes calcium and magnesium precipitate because of high pH.

Function of Dialysate:

1.It carries away the waste materials and fluid removed from the blood by dialysis.

2.It prevents removal of essential electrolytes.

3.It avert excess water removal during dialysis.

Types of dialysate:

Two types of dialysis solutions are discussed in the following 2 tips

1-Acetate Dialysate

2-Bicarbonate Dialysate

## Acetate Dialysate

Acetate is physiologically compatible with blood and metabolized to bicarbonate in the liver. It is mixed with water in proportioning system, usually 1 part concentrate and 34 parts of water, to form the dialysate. A typical composition of acetate containing dialysate (after mixing) contains, Sodium 135-145 mEq/L, potassium 0-4 mEq/L, Calcium2.5-3.5 ,Magnesium 0.5-1 mEq/L, Acetate 35-38 mEq/L, Chloride 100-119 mEq/L, Dextrose 11 mEq/L ,PCo2  0.5 mmHg.

Advantages:

1.Stable during storage

2. Not prone to bacterial contamination

3.A wide variety of formulations are available

4.The delivery system is simple and less costly

Disadvantages:

1.Serum bicarbonate may decrease early during dialysis.

2.Acetate accumulation contribute to cardiovascular instability with vasodilatation and hypotension, nausea , vomiting, and post dialysis fatigue.

3.Acetate dialysate is not suited for high efficiency or high flux dialysis. Serum bicarbonate is depleted and acetate level may exceed the rate at which the liver can metabolize it. Therefore cardiovascular instability may be severe and disequilibrium may occur.

## Bicarbonate Dialysate

Calcium and magnesium will not remain in solution with bicarbonate because of low hydrogen ion content (high pH). To solve this problem, two separate concentrate are used. The proportioning (delivery) system mix and monitor three liquids instead of two.

1.The “A” concentrate

2.The “B” concentrate OR powder.

3. Purified water

      The “A” -indicating acidified- concentrate contains sodium ,calcium , magnesium and potassium, chloride. To maintain low pH enough to keep the calcium and magnesium in solution when mixed into dialysate  a small amount of acetic acid is included.

      The “B” -indicating bicarbonate-concentrate contains the sodium bicarbonate. Sodium chloride may be included in some preparation to raise the total conductivity.

      Dry concentrates:  Manufacturers utilizes a delivery system that accepts a closed container of dry (powder) bicarbonate onto a special holder. Warm water passes  through the column producing a saturated solution of bicarbonate that is proportioned with water then with the A concentrate by conductivity controlled feed back system (e.g. Bicart.). The advantage of this technique avoid the problems of bacterial growth in bicarbonate concentrate and reduces the cost of storage.

- Chemicals in bicarbonate concentrates in mEq/L(after mixing)

| Component | Na | K | Ca | Mg | Cl | Hco3 | CH3COO |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Concentrate A | 81 | 2 | 3.5 | 0.7 | 87.2 | - | 4 |
| Concentrate B | 59 | - | - | - | 20 | 39 | - |
| Total | 140 | 2 | 3.5 | 0.7 | 107.2 | 39 | 4 |

The proportioning system mix the two components simultaneously with purified water to form the product of dialysis solution, just before it goes to dialyzer. During mixing carbon dioxide will be generated as a result acid and bicarbonate reaction. The carbon dioxide generated will form carbonic acid , which lower the pH of the final bicarbonate containing solution to ~ 7-7.4 . In this pH calcium and magnesium remain in solution.

- Negative features of bicarbonate dialysate

1-Liquid “B” is not stable. Some stabilizer or dry sodium bicarbonate may be mixed. The mixing process requires care to avoid much loss of Co2 ; it must be used within 24 hours.

2-Bicarbonate concentrate is very susceptible to bacterial contamination which should be avoided . If the container has been opened for more than 72 hours it should not be used. All containers for mixing, holding, or dispensing B concentrate must be scrupulously sanitized at regular intervals.

3-Many formulation of “A” concentrate are available therefore care must be exercised to ensure that the concentrates selected are correct for the delivery system being used.

Hemodialysis machine consists of  :

1.Blood pump

2.Dialysate delivery system

3.Safety monitors

4.Options

5.System disinfection:

## 1. Blood pump

Function : It circulate the blood through the dialyzer back to the patient.

Kinds    : Presently peristaltic roller pumps which works by progressive compressing special segments of the blood tubing are used.

Pump occlusion: Occlusion means that the roller  compress the tubing segments against the semicircular housing sufficiently to close the lumen completely at that point. Consistent pump flow rates require precise occlusion. Overocclusion may crack the tubing and rupture of the pumping segment. If occlusion is incomplete, there will be backflow of the blood with each stroke, inaccurate flow rate and red cell damage The flow rate for adult is ~ 4 times patient body weight ml/minute, up to 600ml/minute in high efficiency and high-flux dialysis.

## 2.Dialysate delivery system :

Function

-Appropriate blending of concentrate and water for preparation of final dialysate.

-To monitor dialysate for temperature, composition, and blood leak

-To control dialysate pressure or ultrafiltration rate.

-To regulate the dialysate flow rate through the dialyzer.

- Deareation of water

-Provide protective mechanisms to isolate the blood circuit and the patient from unsafe dialysis.

-System for disinfection and cleaning

Types  a. Single batch system    b.Continuous proportioning system     c.Sorbent regenerative system

a. Single batch system     (is of historic interest)

b.Continuous proportioning system (early1960s)

1.Central      :

-All the dialysis solution used for dialysis is produced by mixing the concentrate with purified water and pumped to each machine

-Advantage : Cost effective

-Disadvantage :Not permitting dialysate composition modification

2.Individual    :

-Each dialysis machine proportions its own concentrate with the purified water.

-Advantage :

1. Permitting dialysate composition modification by selecting different concentrate or altering the mixing ratio.

2. There is significant saving in time.

3. Less bacteriologic and chemical contamination. 4. Reduced the risk of human error.

-Disadvantages :

a. Complex sophisticated system

b. Costly

c.Troubleshouting may be difficult and needs a factory-trained service technician.

Methods of mixing in proportioning system

1. Matched pumps

2. Matched metering valves

3. Matched hydraulic-accumulation cylinders

4. Electronic feedback circuit. Details of mixing are beyond the scope the Tips

c.Sorbent regenerative system  for renewing dialysate. At present it is rarely used.

Heating

-The physiological range of dialysate temperature is 36-42 c

-Temperature below 36 patients will complain of cold and uncomfortable.

-Temperature above 42 protein denature can occur and above 45 red blood cells will hemolyze.  cardiorespiratory arrest and death (from hyperkalaemia and hypoxemia )as a consequence of overheated may occur.

Deaeration or Degassing

Water contains considerable amount of dissolved air and microbubbles. Once negative pressure is applied and water is warmed the air comes out of the solution as expanding microbubbles.

Bubbles causes:

1.Artifacts on conductivity, temperature sensors and flow meters

2.When bubbles cross the membrane into the blood foaming occur and increases potential for clot formation, hemolysis and occlusion of the fibers.

3. It decreases the surface area of the dialyzer

4. It block the dialysate flow channels subsequently dialysate pressure drops and the dialysate surface area incompletely utilized for solute transport

Deareation devices uses warmers along with negative pressure (- 600 mmHg) to bring the dissolved air out of the solution.  An air trap or coalescing filter then capture the air and vents them out.

Dialysate pressures

The dialysate delivery system must be capable of generating positive and negative dialysate pressures (~ - 400 to + 200).

         In dialysis machines without an ultrafiltration controller the dialysate pump is located at the line leading from the dialyzer to the drain. This allows the machine to create a negative pressure in the dialysate compartment of the dialyzer. The negative pressure is generated by partial occlusion of dialysate hose proximal to the dialyzer. To increase the rate of ultrafiltration, negative dialysate pressures (suction pressures) are often required. Transmembranous pressure is a resultant of the blood compartment and dialysate compartment pressures.

        High positive pressure is required in situation in which ultrfiltration coefficient or the blood compartment pressure are high to limits the rate of obligatory ultrafiltration. Higher positive pressure (+ 350) may be necessary with high-flux membranes However, backfiltration occurs when dialysate pressure is greater than that of blood compartment, e.g. at the blood outlet. Transport of endotoxin fragments into the blood in such membranes.

Dialysate Flow

The standard dialysate flow rate is 500 ml/minute.  When high-efficiency and high-flux  dialyzer are used high dialysate flow rate between 700 and 1000 is required along with high blood flow rate.   High flow rates are needed  for:

1. Optimum use of dialyzer surface area for solute transport.

2.With high dialysate flow rate a positive dialysate pressure is generated which limits the rate of obligatory ultrafiltration.

3. Safety monitors:

Monitors for blood circuit

1.Arterial pressure monitor

Locations:  Proximal to the blood pump

Function:  1. It reads the arterial pressure at the segment between the patient's needle site             and proximal to the blood pump which represent the negative pressure created by the roller pump. 2.It identifies how much suction is being placed on the arterial wall and guard against excessive suction on the vascular access, e.g. .if suddenly the arterial pressure increases from 100 to 200 mmHg this could indicate clotted or dislodged needle, low patient BP or kink in the arterial line. 3. Resistance within the needle (function of the gauge and needle length). 4. It provide an index of vascular access blood supply relative to the flow demand by the blood pump. 5. A guide to appropriatence of needle placement or kink or obstruction in the blood segment between the patient and the monitor.

Mechanism: The pressure is  monitored by mechanical or electronic manometers (pressure transducers) . Electronic transducer is more sensitive and have rapid response

2.Venous pressure monitor

Location: Just distal to the dialyzer, usually attached to the top of the venous air trap.

Function: 1. It reads the venous pressure at the segment between the point after the dialyzer and before the blood reenters the patient's body. 2. It represent the resistance of the blood returning to the patient via the venous needle e.g. if suddenly the venous pressure increases from 50 to 150 mmHg this could indicate kink in the venous line, clotted air trap or the venous  needle may be clotted or maligned. Sudden dropping venous pressure occurs when the venous needle is pulled out, wet transducer or clotted arterial chamber.3. It may indicate venous stenosis proximal to the needle site.

Mechanism:  As in arterial pressure monitor.

3.Venous air trap and air detector

Location: 1. Just distal to venous pressure monitor.

2. Often second air trap on the arterial line is also used.

Function: To prevent air entry into the patient or to the dialyzer. It is used also to measure the pressure in that segment of blood circuit.

Mechanism: When air bubbles entered the blood circuit a sensor reacted through ultrasonic transducer or light beam by stopping the blood pump, clamping the venous line, and activating audiovisual alarms.

Monitors for dialysis solution circuit

1.Conductivity

Location :   Before dialysate reaches the dialyzer

Function    To guard against excessive diluted or concentrated dialysis solution.

Mechanism : When conductivity meter (conductivity probe) detects high or low conductivity  the machine automatically sounds an alarm and puts the dialysate into bypass mode so that  no dialysate flow to the dialyzer.

Exposure of the blood to hyperosmolar dialysate can lead to hypenatraemia and other electrolyte disturbances. Exposure of the blood to hypoosmolar dialysate can results in  hyponatraemia and rapid hemolysis.

2.Temperature

Location:    Before the dialyzer

Function:    To avoid high temperature

Mechanism: Through a temperature sensor, high temperature activate an audiovisual alarm simultaneously with bypass mechanism.

3.Bypass  valve or mechanism

If the conductivity or Temperature found to be out of limits a bypass valve is activated to divert dialysate around the dialyzer directly to the drain.

4.Blood leak detector

Location : In the effluent dialysate line

Function: It guards against undetected blood loss during dialysis. The maximum limit to hemoglobin present in dialysate is 70 mg/L (corresponding to blood loss of ~0.4 ml/minute ) after which the blood pump is stopped.

Mechanism: A beam of light is directed through a column of dialysate onto a photoelectric cell.  A change in translucence and light scatter in dialysate reduces the light received by the photocell, stopping blood pump and activating audiovisual alarms.

5.Dialysate pressure or transmembranous monitor

Function : To monitor the ultrafiltration with an upper limits avoids excessive level (the usual preset range 0 to -500 mmHg). Excessive TMP may lead to rupture of the dialyzer and secondary deareation (cause air accumulation).

Mechanism: The monitor may have automatic or manually set limits, so that extrusion outside the limits trigger an audiovisual alarm.

4. Options in hemodialysis machine:

1.Heparin pump

2.Bicarbonate  3.Variable sodium

4.Controlled Ultrafiltration

5.Programable ultrafiltration

6 Dialysate urea sensor [on-line Kt/V monitor] 7.Single blood pathway.

## 5. System disinfection:

All dialysis unit s must have written policies the deal with the dialysis fluid pathway and dialysis machine . Disinfection procedure should be done on regular base according to manufacturer's instructions.

Target: To control bacterial contamination. HIV, HCV and HBV viruses are known to be inactivated by common household bleach.

Methods

1.Heat disinfection requires temperature greater than 85-90 C

2.Chemicals disinfection such as formaldehyde ,  sodium hydrochloride and acetic acid

Machine surfaces, patient's chairs, surrounding furniture, equipment should be routinely wiped with 10% bleach solution following every patient dialysis. Blood spills should be cleaned immediately.  Leak proof bags should be used to transport linen soiled with blood or body fluid

## 6. Alarm during Hemodialysis- Problem Solving Guide Lines]

When an alarm is activated during dialysis do the following:

1-Identify which alarm has been activated.

2-Identify the cause.

3-correct the cause

4-Resume dialysis if safe to do so

| Problem | Management |
| --- | --- |
| 1.Power | a. Turn the system off and on; if still no power, check the power cord; make sure power is available b. Check the fuse |
| 2. Arterial and venous Pressure | a. Check to see that blood pump is running and connected properly b. Check to see if blood flow rate has changed c. Determine if patient has coughed or moved d. Check to see if the monitor line is leaking e. Check the blood line for kinks or leaks f. Ensure the monitoring lines are connected to proper drip chambers |
| 2.1 High venous pressure | 1-Manipulate the needle and or the line. If the access is small, a tourniquet must be used, being certain the blood pump is off ..Recantation with new needle if needed ,the original one should be left in place until the end of dialysis. to avoid undue bleeding as the patient is heparinized 2-Adjust the blood flow rate , Proper heparnization , Treat access problem ,and Proper needle and needling 3-Extreme care must be exercised when dialyzing a patient with high venous pressure -This increases the baseline TMP and obligatory ultrafiltration will occur -Single-needle device is occasionally impossible to use, because with high venous pressure ,venous return will be impaired and blood recirculation will be high |
| 2.2 Increase in negative pressure: (Excessive inflow suction) | Management of increased negative pressure: 1- Manipulation of the arterial needle is similar to that of venous needle. In most cases the needle may have to be replaced with the same precaution as with venous needle 2-Treat the cause. 3- Proper needling 4-Asses the access |
| 3.Air detector alarm | a-Check for air leaks around tubing joints. b-Excessive undetected negative pressure c-Check for unattended intravenous solution administration. Management of air embolism if present a-Clamp the venous line and stop the blood pump. b-Place the patient in Trendelenberg position on the left side with the chest and head tilted downward to trap the air at the apex of right ventricle away from the outflow tract. c-Cardiorespiratory support ,Oxygen 100%. d-Occasionally percutaneous aspiration of the air foam from the heart may be necessary. e-Other measures include IV Dexamthsone to reduce brain edema ,Heparin /Dextran to improve the microcirculation. |
| 4.High Positive dialysate pressure alarm | a Check to see if drain is occluded or kinked b. Check to see if dialysate hoses are leaking air |
| 5.Low Negative dialysate pressure alarm | Check to see if dialysate hoses are kinked |
| 6.HighTemperature alarm | a. Determine temperature of dialysate (to dialyzer) in the line actually high b. Check to ensure incoming water temperature is blew 90 F |
| 7.Low Temperature alarm | a. Determine temperature of dialysate”to dialyzer” in the line is actually Low b. Turn the mode selector switch to “dialyze” c. Allow adequate time for the system to stabilize and come to proper temperature range d. Check to ensure incoming water temperature is above 40 F |
| 8.High conductivity alarm | a. Check to see if water is flowing too slowly or turn off b. Check for kink in the concentrate out line c. Make sure that the system has had time to stabilize d. Analyze the dialysate to confirm high conductivity at the “to dialyzer” line connection : 1.If normal, there is malfunction in the machine itself (change it) 2.If high again concentrate should be changed 3.Check the conductivity before resuming dialysis e. Be certain that the dialysate flow rate is proper. |
| 9.Low Conductivity alarm | a. Turn the mode selector switch to “dialyze” b. Connect the concentrate line to the system c. Drop the concentrate line into the concentrate container d. Check for kinks in the “concentrate” in line e. Make sure the filter on the “concentrate in” is clean g. Allow adequate time for the system to stabilize h. Change the concentrate container if it is dry i. If the concentrate container has not run dry ,a sample of dialysate should sent to the laboratory for sodium and chloride j. Analyze the dialysate to confirm low conductivity at drain 1.If low conductivity is confirmed ; change the concentrate bottle (recheck again) 2.If the conductivity still low after changing the concentrate , a different machine should be tried l. Be certain that the dialysate flow rate is proper |
| 10.Blood Leak alarm | a. Make sure the blood leak detector is clean b. Check the effluent for traces of blood c .Check the dialysate lines for gross blood leak d. Exclude 1.bilirubin in dialysate in jaundiced patient 2.air bubbles in dialysate 3.dirty sensor e. If a leak confirmed and you can not mange the cause 1.reduce dialysate compartment pressure to -50 mmHg to avoid bacterial entery to the blood 2.Dicontinue dialysis |

Hemodialysis Procedure

Schematic representation of dialysis system

## 1. General Reassessment

Acute or chronic dialysis prescription should be reviewed. evaluated. and carried out accurately to obtain the maximal efficiency for dialysis. The patient's physiologic status is assessed to ascertain the necessity of adjusting any dialysis orders.  All machine parameters are assessed to ensure that the prescribed procedure is correctly implemented. The goal is to initiate and terminate the dialysis procedure safely and comfortably with no or minimal complications.

## 2.Rinsing and priming

Technical and nursing personnel have a great responsibility to assure proper and safe connections, flow paths, and overall adequacy and safety of the system.

Adherence to the manufacturer's procedure for rinsing is mandatory.

Importance of rinsing and priming :

1Rinsing and priming the dialyzer assure adequate removal of allergens (e.g. ethylene oxide ) and reduction of the incidence of anaphylactic membrane reactions

2.Microbubbles are also removed when the venous end of the dialyzer pointed upward.

The dialyzer should be used within 5-10 minutes to avoid leaching of residual ethylene oxide or other leachable allergens into the rinsing fluid.

Dialyzer should be rerinsed briefly immediately prior to dialysis if more than 10 minutes has elapsed.

3. Patient Monitoring Predialysis:  (See its Tip)

Weight , Pulse Rate , B.P.Laying   & Standing  Temperature, Fluid status, Blood investigations and Vascular access patency and freedom from infection

## 4.Obtaining vascular access

Poor vascular access is a limiting factor to patient survival on hemodilysis. Therefore great care must be taken to maintain adequate vascular access.

        Percutaneous venous cannula ( Femoral,  Subclavian,  Jugular )

Residual heparin or clot is first aspirated from both catheter lumen

Check the patency of each lumen by irrigating with a saline filled syringe.

Heparin loading dose is administered in the venous limb and flushed with saline.

Initiate dialysis after 3 minutes

        AV fistula and graft: (see the previous tip; Using permanent vascular access)

Heparin loading dose is administered in the venous needle and flushed with saline.

Initiate dialysis after 3 minutes

## 5.Initiatting dialysis

        Set the blood flow rate at 50 then 100 ml/minute, untill the blood fills the blood circuit.

        The Priming fluid in the lines and dialyzer is disposed of to drain until the blood reaches the venous air trap. In unstable patient the priming fluid is usually given to the patient t maintain the blood volume.

        Increase the blood flow rate to the desired level after the circuit is filled with blood (150-250 in acute cases).

        Initiate the dialysis solution flow and adjust the TMP.

6.Alarms ( See the previous tip )

Blood circuit

Arterial pressure

Venous pressure

Air detector

Dialysis solution circuit

Conductivity

Temperature

Blood leak

7. Patient monitoring During Dialysis (see its Tip)

Pulse rate , BP every 30 to 60 minutes in chronic dialysis, but at least every 15 minutes in acute dialysis, Food & Fluid intake, Complications during dialysis and any particular observations.

## 8. Termination of dialysis

a. Saline rinse: The blood is returned by pumping sterile normal saline into the arterial side until the blood is displaced.  After the bubble trap the fluid should be very pale pink in color (to assure that the patient has lost the least amount of red cells).

b. Saline-Air rinse: The blood is forced by pumping a small amount of saline into the arterial line, then the line is opened to allow air into the circuit to push the saline and blood. Again the fluid entering the patient should be very pale pink in color.

c. Air-Saline rinse: It begins with an infusion of air to displace the blood until it reaches the end of the dialyzer. At this point the saline is pumped into the arterial line to displace the air.  At the same time the air is allowed to escape at the level of venous chamber while the saline is being infused to the patient. It is effective, but air embolism is a potential risk.

NB: I- Constant visual monitoring of venous line is required to avoid air infusion into the

patient in b and c methods.

II- Hollow fiber dialyzer are rinsed by saline rinse method because complete removal

of blood by air from tiny diameter of hollow fiber is difficult.

9. Patient monitoring Postdialysis    (see its tip)

Weight  ,  Pulse Rate  , BP Laying &Standing ,Temperature, Blood investigations and Vascular access patency. All patients parameters and any unusual occurrences should be documented on patient's file.

## 10. Equipment care

The care of dialysis machine is the responsibility of the staff and  of the biomedical technicians. Scheduled maintenance recommended by the manufacturer should followed meticulously for the safe and efficient function of the equipment.

Anticoagulation for Hemodialysis

Mechanism of clotting in extracorporeal circuit:

Clotting of blood in extracorporeal circuit may be initiated when blood comes into  contact with cannulas , lines, and the dialysis membranes.

Steps of clotting: 1.Coating of the circuit with plasma proteins.

2.Platelet adherence and aggregation

3.Generation of Thromoxane A2.

4.Activation of intrinsic coagulation cascade.

5.Thrombin formation and fibrin deposition.

Predisposing factors:

1.Reduced anti coagulant.

2.Reduced blood flow rate.

3.Increased whole blood hematocrit [EPO therapy]

4.Increased extracorporeal hematocrit [excessive ultrafiltration].

5.Blood transfusion and Lipid infusion during dialysis..

Signs of blood clotting in the extracorporeal circuit :

1.Dark-colored blood

2.Dialyzer :-Presence of black streaks in the dialyzer

-Reduced residual dialyzer  volume

-Presence of clots at arterial header.

3.Lines -Foaming at the drip chambers and venous trap.

-Clot formation at the drip chambers and venous trap.

4.Pressure :-Increased difference between the postpump and venous pressure

when there is clots at arterial chamber or the dialyzer.

-Increase in postpump followed by increase in venous pressure

if the clots are distal to the venous chamber

-Increased venous pressure when venous needle is clotted

Anticoagulants

1- Heparin       a.Crude heparin   b.Low molecular weight heparin

2 Antiplatelet agents

a)Protaglandins PGI 2 , PGE 2 ,eporostemol and iloprost

-Potent inhibitors of platelets aggregation

-Less effective than heparin

-Side effects include ; hypotension ,flushing, headache ,nausea and vomiting

b)Aspirin , NSAID , Sulphinpyrazone and Ticlopidine

-Antiplatelets + Counteract platelet factor 4 and prevent its heparin neutralizing effects

-Not suitable to maintain anticoagulation -May be used as heparin sparing

.    3- Protease inhibitors  [Nofomostat, Gabexate]

-Inhibitors for both coagulation/fibrinolysis cascade and platelet aggregation.

-Adequate anticoagulant effect and reduce bleeding complications.

4-Hirudin

-It is polypeptide thrombin inhibitor.Unlike heparin does not require cofactor to act

-Does not cause platelet stimulation or aggregation

- Heparin and Heparinization

Nature of Heparin

Heparin is an anionic sulfated mucopolysaccharide of variable molecular weight [8000 to 14000 d] .Low Molecular Weight Heparin fractions [4000-6000 d] are obtained from crude Heparin .  Heparin is strongly acidic And is neutralized by strong basic compounds such as protamine ,toluidine and quinidine . The half life is about 90 minutes and the peak anticoagulant activity is reached 5-10 minutes

Mechanism of action:

Heparin alone does not have anticoagulant effect. In the blood it combine with a protein fraction called heparin cofactor [antithrombin III].

The complex of Heparin -Antithrombin III prevents  clotting at the three stages of coagulation 1-It binds and inactivate Thrombin

2- It binds and inactivate Activated factor X

3-It binds and inactivate Activated factor XI

Low Molecular Weight Heparin [LMWH] inhibits coagulation Activated factor X , XII and kallikrin, but little inhibition on Thrombin, factor X and XI therefore PTT and thrombin time are minimally prolonged,thus reducing bleeding risk.

Side effects :

Bleeding             Pruritus                Allergy

Osteoporosis      Hyperlipidemia     Thrombocytopenia

Monitoring clotting times during hemodialysis

Several laboratory method are used:

1-Activated partial thromboplastin time [APTT or PTT]

2-Activated clotting time [ACT]

3-Lee White clotting time [LWCT]

Heparinization  \   Test                                          PTT              ACT              LWCT

Baseline                        60-85 sec        120-150 sec      4-8 min

Standard      - During dialysis               120-140          200-250            20-30

-At end of dialysis             85-105           170-190              9-16

Tight            -During dialysis                 85-105            170-190              9-16

-At end of dialysis             85-105            170-190              9-16

The goal  is to maintain the PTT or ACT at the baseline value plus 80% during dialysis and plus 40% only at the end of dialysis to minimize bleeding risk after withdrawal of access needles.

Techniques of heparin administration

1-Intermittent infusion

An intravenous loading dose of heparin [25-50 units/kg of body weight] is given via the last needle placed ,and smaller maintenance doses are repeated intermittently throughout the procedure. Clotting times are monitored at intervals and the dose is adjusted accordingly.

2-Continuous infusion

The patient is given the loading dose followed by slowly infusion of the maintenance dose to the blood circuit at a constant rate throughout the dialysis .Clotting times are done at intervals and the rate is adjusted accordingly.

Remarks

1-In continuous infusion the loading dose of heparin is lower than with intermittent .In the former the dose is required only to prolong the clotting times to the baseline plus 80%  while in the later it is required to prolong the clotting times to above the base line plus 80%.

2-The dose of heparin is depending on sensitivity to heparin of the individual patient ,heparin half life and potency of heparin preparation

3-For a patient with an average heparin half life of 1 hour, stopping heparin administration approximately 1 hour prior to the end of dialysis .

4-During heparinization tendency to bleed is potentated by uremic platelet dysfunction and by endothelial abnormalities.

3- Regional heparinization:

Administration of heparin “regionally” into the arterial line blood  and neutralizing it by administration of protamin sulfate into dialyzed blood before returns to the patient.

Protamin is strongly basic compound acts by combining chemically with the strongly acid heparin to form a stable complex with no anticoagulant effect.

Dissociation of heparin-protamin complex may occur up to 10 hour post dialysis causing bleeding problem, this is called heparin rebound. Flushing ,bradycardia ,hypotension ,and dyspnea are other side effects of protamin .

4-Tight heparinization:

Tight or low dose heparinization is indicated for patient at slight to moderate risk for bleeding e.g. pericarditis and recent surgery

Loading dose 10 units/kg  followed by small additional doses according to the target clotting times. Continuous administration is preferable in tight heparinization to avoid fluctuation in clotting during dialysis.

- Hemodialysis without heparin [heparin free dialysis] :

Indications In patients with high bleeding risk

1.Pericarditis

2.Recent surgery with bleeding complications or it will be dangerous e.g. cardiac

vascular , eye ,renal transplant and brain surgery

3.Blood disease ; coagulopathy , thrombocytopenia

4.Intracerebral hemorrhage

5.Any active bleeding

Methods of heparin free dialysis

1.Heparin rinse

a,Rinse the extracorporeal blood circuit with 3000 units heparin/liter

b.Flush the heparin containing saline with unheparinized saline or patient blood.

2.High blood flow rate

Set the blood flow rate as high as possible , 250-300 ml/minute if it can be tolerated.

3.Saline rinse Rinse the dialyzer with 100-200 ml of saline .Adjust the UF rate to remove the excess fluid

Patient monitoring is a series of repeated or continuous observation of the patient’s appearance and physiologic state before, during and after dialysis procedure. These observations are charted and made part of the patient’s permanent record. The objectives are to provide as comfortable and safe procedure as possible for the patient and to detect as early as possible  any complication during dialysis.

General Observations

Monitoring Predialysis:        Weight , Pulse Rate , B.P.Laying &Standing  , Temperature

Monitoring During  Dialysis: Pulse Rate , BP , Food &Fluid intake

Particular Observations

Monitoring Postdialysis:      Weight  ,  Pulse Rate  , BP Laying &Standing ,Temperature

.

Weight:

The weight of the patient is a good index of how well , or how poorly , the patient is controlling his fluid balance between dialyses. The pre and postdialysis weight provide the best indication of the amount of  ultrafiltration needed during the procedure. The patient should be weighted immediately before and after each dialysis, wearing the same articles of clothing and using the same scale. the patient should weigh himself or herself daily. One should strive to keep the interdialysis weight gain below  1.0 kg/day

The dry weight is the target postdialysis weight that ideally would results in removal the excess body fluid.. The dry weight for each patient must be determined on trial-and-error basis. If the dry weight is set too high the patient will have clinical evidence of fluid overload. On the other hand if it set too low the patient may suffer from malaise ,dizziness ,weakness , cramps , and frequent hypotension episodes.

Pulse Rate:

At the start of dialysis, pulse , temperature, BP observations serve as baseline A rapid pulse may indicate low hematocrit or fluid overload. An increase in pulse rate during dialysis my be associated with  decreasing blood volume from ultrafiltration and may occur before blood pressure drop. Arrhythmia may indicate some complications e.g. ,cardiac instability, electrolyte disturbance ......etc.,it should be brought to the physician’s attention.

Temperature:

High temperature suggests infection or complicating illness. The patient should be questioned as to other symptoms The vascular access should be inspected carefully for evidence of infection, swabs and cultures should be collected ,and evaluation by the physician is mandatory. Temperature during dialysis may be the result of warm dialysis fluid , a pyrogen reaction or  showering of the bacteria from the infected access.

Blood Pressure:

      Blood pressure is the pressure exerted by the blood against the walls of the arterial blood vessels during systole and diastole of the heart. It results from the pumping force of the heart and the resistance of the vessels.

Usually BP is measured indirectly with a cuff-type sphygmomanometer. Occasionally both arms are used to complete the blood circuit, in such instance, the cuff can be wrapped  around  the midthigh area and the stethoscope applied at the bend of the knee. The BP obtained by this method, however, will be 20-40 mm Hg higher than arm pressure.

      For patients on dialysis, normal BP is largely individual matter. That is, we are interested in changes that may occur than in absolute values. The patient’s BP should be monitored every 30 minutes for an acute dialysis and every 30-60 minutes for chronic dialysis. A systolic value greater than  200 or diastolic greater than 110 should be brought  to   physician’s     attention. Predialysis hypertension is usually  volume  related. Lowering dialysate sodium concentration (not below 135-140 mEq/L) and restriction of salt fluid intake is of great help to control predialysis hypertension. Most of the patient are instructed not to take antihypertensive medication prior to dialysis .In severely hypertensive patient , patients being dialyzed in the afternoon and those patient if hypotension during dialysis is not a problem., they can take their antihypertensive medication safely.

Causes ,prevention ,and management  of interdialytic hypotension were discussed previously.

Food and Oral Fluid intake:

      There are several reasons for watching food and oral fluid intake during dialysis. The amount of fluid removed by ultrafiltration  is estimated by the net change in the weight from the predialysis to the postdialysis state. The quantity of food or fluid ingested should be taken into consideration in making this calculation. A pint of fluid is equals to half kilogram. Although it is permissible for most patients to eat during dialysis ,if desired ,it is best to limit this to a small meal or snack.

      Digestion may contribute to development of hypotension and vomiting. This is distressing and increases the risk of aspiration as well as interfering with a smooth dialysis.

      Ice chips can be used as water substitute, it is effective  in alleviating the sensation of thirst than water. The disadvantage of ice is that is water, and people tend to disregard this fact. A 200 ml of ice chips is equivalent to 150 ml of water.

     It is not recommended that to oral fluids be given in large amount as a control  for excess ultafiltration. Such fluids may reach the intracellular compartment during dialysis and then be difficult to remove. Also, oral fluid is always hypotonic, contributing to hyponatremia.

.

Particular Observations:

General condition  and response of the patient during the procedure .Nausea ,Vomiting ,Apprehension, Shortness of breath, Chest pain, Sweating, Pallor, Restlessness or agitation , Irritability, Itching, Flushing , Childish or Hysterical behavior, Sleepiness ,and complaints of pain are some of many points to be noted.

.

Laboratory tests

The plasma urea nitrogen and s creatinine are monitored monthly. The midweek, predialysis level is commonly followed. The plasma urea level is largely determined by the amount of protein ingested, while the plasma creatinine level depends on the muscle mass. Factors other than protein can affect plasma urea .If the plasma urea is higher than expected many factors should be excluded e.g. Increased dietary intake, Hypercatbbolic state , GIT bleeding , V.Access Recirculation, Dehydration. On the other hand if it is lower than expected; Decreased dietary protein intake, Increased dialysis treatment, and Liver diseases should be excluded. For patients with high predialysis  S.Potassium diet control and resin exchange is necessary.

      Monitoring S Calcium, S Phosphate and Alk Phophatase are helpful tests to prevent and treat renal bone disease These values are usually checked monthly predialysis. The target  plasma calcium should be at the upper level of normal and plasma phosphate should be at the lower level of normal.

LFTs are usually checked monthly ,and may unmask silent liver disease ,especially Hepatitis or Hemosidrosis. The aim of monitoring Hemogram ,S Iron ,Tranfserrin ,Ferritin is to detect and treat Hematological Abnormalities in dialysis patient

| Frequency | Test |
| --- | --- |
| Monthly | 1-Renal profile (Urea ,Creatinine, Electrolyte) 2-S.Calcium ,Phosphate ,Alk Phosphatase 3-LFTs(TSP-S Albumin, AST, ALT ,S Bilirubin) 4-Blood Glucose 5-Hemogram (WBCs ,Hb, PCV, MCV, PLAT.,) 6-S Iron ,Transferrin 7-KT/V |
| Every 3 months | 1-HBsAg,Anti HBsAg 2-Anti HCV 3-HIV 4-S.Aluminum |
| Every 6 months | 1-S.Ferritin 2-PTH |
| Yearly | 1-Skletal Survey 2-X-Ray Chest 3-ECG |

Common Complications during Hemodialysis

1- Hypotension (20-30% of dialyses)

2- Muscle Cramps (5-20% of dialyses)

3- Nausea  and Vomiting (5-15%of dialyses)

4- Headache (5% of dialyses)

5- Chest Pain (2-5% of dialyses)

6- Back Pain (2-5%of dialyses)

7- Itching (5% of dialyses)

8- Fever and chills (<1%of dialyses)

Serious complications during hemodialysis

1-Disequilibrium syndrome

2-First use syndrome

3-Arrhythmia

4-Cardiac tamponade

5-Intracranial bleeding

6-Seizures

7-Hemolysis

8-Air embolism

.

## b. Mechanical Complications

1.Rupture dialyzer

2.Clotted dialyzer

3. Air embolism

4.High conductivity

5.Low conductivity

6.Low water pressure

7.High venous pressure

8.Abnormal arterial pressure

9.Electrical failure

10.Needle-site bleeding

11.Hemolysis

## Management of Medical and Mechanical Complications During Hemodialysis

| Complication | Management |
| --- | --- |
| Hypotension | 1-Place the patient in trendlenberg position (if respiratory status allows). 2-A bolus of 0.9 saline (100-250 ml) should be administered through the venous line. Alternative to 0,9 saline : -3%saline 50-100 ml -Dextrose 50 % 25-50 ml. -Human Albumin 20 % 50-100 ml. -Mannitol. -Pressor agents egg Dopamine. 3-Reduce the UF rate near to zero. 4-Discontinue dialysis in severe cases CLICK HERE FOR MORE DETAILS |
| Disequilibrium syndrome | In mild cases : Treat the symptoms , Reduce blood flow rate, Hypertonic saline dextrose solution can be administered. In severe cases: Stop dialysis ,treat the seizures Supportive measures of coma , Mechanical ventilation if necessary, IV mannitol may be of benefit. CLICK HERE FOR MORE DETAILS |
| Angina | 1.Nasal oxygen should be initiated. 2.Reduce blood flow rate and stop ultrafiltration. 3.Sublingual nitroglycerin if blood pressure is maintained. 4.Sedation. 5 Treat the possible cause. 6.Dialysis with bicarbonate dialysate bath. 7.In severe cases discontinue the procedure . 8.ECG and Further investigation may be required to exclude myocardial infarction. CLICK HERE FOR MORE DETAILS |
| Air Embolism | a-Clamp the venous line and stop the blood pump. b-Place the patient in Trendelenberg position on the left side with the chest and head tilted downward to trap the air at the apex of right ventricle away from the outflow tract. c-Cardiorespiratory support , Oxygen 100%. d-Occasionally percutaneous aspiration of the air foam from the heart may be necessary e-Other measures include IV Dexamthsone to reduce brain edema ,Heparin /Dextran to improve the microcirculation. CLICK HERE FOR MORE DETAILS |
| Seizures | 1. Discontinue dialysis. 2.Ensure airway patency 3.Collect blood sample for glucose, calcium, electrolyte, urea and creatinine. 4.Administration of 5-10 mg Diazepam slowly IV. Repeat after 5 minute intervals, if seizures persist, to a maximum 30 mg. 5.Give glucose IV if hypoglycemia is suspected. 6.Give calcium gluconate IV if hypocalcaemia is suspected. 7.Treat other electrolyte disturbance. |
| Muscle Cramps | 1-Concomitant occurrence of hypotension and cramps may respond to treatment with 100ml of 0.9 % saline 2-Hypertonic saline or dextrose [ Leads to dilatation of the muscle bed blood vessels restore blood volume]. *Hypertonic dextrose( 50 ml of 10 % dextrose.) is preferred for treating non diabetic patient to avoid thirst produced by saline. CLICK HERE FOR MORE DETAILS |
| Nausea and Vomiting | 1.Reduce blood flow rate if acetate dialysate is being used by 30 % during first hour of dialysis 2.Reduce ultrafiltration rate 3.Metoclopramide Hcl [Primperan] 10 mg IV, or Prochlorperazine [Stemetil]12.5-25 mg IV 4.Treat the cause / Use bicarbonate containing dialysate |
| Headache | 1.Reduce blood flow rate if acetate dialysate is being used by 30 % during first hour of dialysis 2.Acetaminophen tab [Paracetamol] 500 -1000 mg PO during dialysis. 3.Treat the cause / Use bicarbonate containing dialysate |
| Itching | a-Drug therapy : -Antihistamines,e.g. Diphenhydramine ,Hydroxyzine -Parenteral lidocaine infusion -Oral activated charcoal b- General measures. 1- Relieve the dry skin with topical emollients e.g. Lubricating solutions , Camphor and Menthol containing creams 2-Switch from ethylene oxide to gamma ray-sterilized dialyzer. c-Phototherapy: Ultraviolet light-B delivered in a conventional light box giving 2 treatments per week for 4 weeks.. d-Control of calcium and phosphorous metabolism: CLICK HERE FOR MORE DETAILS |
| Needle-site bleeding | 1-Direct pressure is the simplest and effective measure. It should be aseptically and with extreme care to avoid AV access occlusion. 2- Apply gelfoam to the puncture site 3-Adjust heparin dose CLICK HERE FOR MORE DETAILS |
| Hemolysis | 1-Clamp the blood lines 2-Turn the blood pump off 3-Do not reinfuse the blood 4-Leave the needle in place 5-Draw blood for: Type and cross match CBC Electrolyte Trace metals, 6-Save dialysate and blood samples for further evaluation 7-Discard the dialyzer line 8-Evaluate the problem by reviewing the following a. Check conductivity b.Check the dialysate temperature c.Look for kinks in the blood line in the pump d.Check dialysate for ;Electrolyte Formaldehyde-Chlorine-Trace metal 9-Check the IV priming solution !0-Resume dialysis to prevent hyperkalemia and blood transfusion if necessary CLICK HERE FOR MORE DETAILS |
| Ruptured Dialyzer | An immediate decision must be made to: 1- Reinfuse the blood (or not) 2-Change the ruptured dialyzer as quickly as possible. 3-Resume dialysis if safe to do so [Usually less than 1% of blood present in dialysate trigger the blood leak alarm] CLICK HERE FOR MORE DETAILS |
| Clotted dialyzer | 1.When a dialyzer clots , the dialyzer ,the arterial line, the venous line may need to be replaced. 2.In clotted dialyzer without rupture , the entire dialysate circuit does not need to be set up again or recleaned 3.Correct the cause CLICK HERE FOR MORE DETAILS |
| High Venous Pressure | 1-Manipulate the needle and or the line. If the access is small, a tourniquet must be used, being certain the blood pump is off.. Recannulation with new needle if needed, the original one should be left in place until the end of dialysis. to avoid undue bleeding as the patient is heparinized 2-Adjust the blood flow rate, Proper heparnization, Treat access problem, and Proper needle and needling 3-Extreme care must be exercised when dialyzing a patient with high venous pressure *This increases the baseline TMP and obligatory ultrafiltration will occur *Single-needle device is occasionally impossible to use, because with high venous pressure ,venous return will be impaired and blood recirculation will be high. |
| Abnormal Arterial Pressure | 1- Manipulation of the arterial needle is similar to that of venous needle. In most cases the needle may have to be replaced with the same precaution as with venous needle 2-Treat the cause. 3- Proper needling 4-Asses the access CLICK HERE FOR MORE DETAILS |
| High Conductivity | 1.Evaluate the incoming water and check for kink the line. 2.The water pressure and filters should be also checked 3.Recheck the conductivity: If normal, there is malfunction in the machine itself(change it) * If high again concentrate should be changed * Check the conductivity before resuming dialysis CLICK HERE FOR MORE DETAILS |
| Low Conductivity | 1.Change the concentrate container if it is dry 2.If the concentrate container has not run dry, a sample of dialysate should sent to the laboratory for sodium and chloride 3.Recheck the conductivity with second conductivity meter: *If low conductivity is confirmed; change the concentrate bottle (recheck again) *If the conductivity still low after changing the concentrate, a different machine should be tried CLICK HERE FOR MORE DETAILS |
| Membrane Reactions | Type A 1-Stop dialysis 2-Clamp the blood lines 3-Do not reinfuse the blood 4-Discard the dialyzer and the line 5-Cardiorespiratory resuscitation 6-Antihistaminic ,Steroids ,Epinephrine can be given Type B 1-Oxygen therapy 2-Treat as in angina during dialysis CLICK HERE FOR MORE DETAILS |

.

## Disequilibrium syndrome

Definition: It is a set of systemic & neurological symptoms and EEG finding that occur during or soon after acute  dialysis

Pathogenesis:

1-Rappid removal of solute from the extracellular fluid  compartment  results  in an osmolar gradient between brain and blood  hence results in cerebral edema.

2-A fall in pH of spinal fluid has been also noted to occur.

Clinical manifestations :

| Mild | Severe |
| --- | --- |
| Headache Fatigue Nausea Vomiting Muscle cramps Restlessness | Hypertension Agitation Confusion Seizures Stupor Coma |

Prevention:

1-Limit the amount and rate of dialysis.

2-Use  of high dialysate sodium levels.

3-Use bicarbonate dialysis.

4-Intavenous infusion of:

Hypertonic dextrose

Hypertonic saline

Mannitol

Management:

In mild cases :

Treat the symptoms,

Reduce blood flow rate,

Hypertonic saline or dextrose solution can be administered.

In severe cases:

Stop dialysis, treat the seizures

Supportive measures of coma

Mechanical ventilation if necessary,

IV mannitol may be of benefit.

N.B. If coma is due to disequilibrium, the patient should improve within  24 hours.

.

## Intradialytic Hypotension

Causes

A-Decreased Plasma Volume :

1-High ultrfiltration rate

2-Fluctuation in the ultrafiltration rate.

3-Target dry weight set low

4-Low dialysate  sodium.

5-Dialyzers with large surface area

6-Increased blood flow rate.

7-Increased dialysate flow rate

B-Decreased Compensatory Related Vaso Constriction:

1-Acetate dialysate.

2-Dialysate temperature 37-38’c.

3-Low dial ysate calcium.

4-Antihypertensive medications.

5-Biocompatible dialyzer membrane

C-Cardiac Related Factors:

1-Failure to increase cardiac rate :

Ingestion of B blockers.

Uremic autonomic neuropathy.

Aging

2-Inability to increase cardiac output :

Poor myocardial contractility due to aging.

Hypertension.

Atherosclerosis.

Myocardial calcification.

Valve disease.

Amylodosis.

D-Other Causes:

1-Pericardial  temponade.

2-Myocardial infarction.

3-Occult hemorrhage.

4-Septicemia.

5-Arrhythmia.

6-Anaphylaxis.

7-Hemolysis.

8-Air embolism.

Management of Intradialytic Hypotension

A-Measures to Ameliorate :

1-Accurate evaluation of dry weight ; is the most important step in minimizing intradialytic hypotension . limit weight gain to less than 1 kg / day .

2-Dosing of antihypertensive drugs . Give the daily dose after dialysis ( not before ) .

3-Adequate dialysate sodium concentration (or high sodium) has been effective in reducing intradialytic hypotension.

4-Bicarbonate dialysis & avoid acetate dialysis in those patient.

5-Cooled dialysate ; in some patient ; dialysate temperature of 35.5’c may be adequate to overcome the predisposition to hypotension.

6-Sequential UF is less likely to cause hypotension.

7-Oxygen supplementation during  dialysis specially in elderly is valuable to overcome predisposition to hypotension (to avoid hypoxemia).

8-Dialysate calcium concentration of 7.5 mg /dl may prevent hypotension

9-Infusion of Human Albumin before dialysis is also helpful.

10-The use of biocompatible dialyser decrease the incidence of hypotension

11-Placement of pressurized leg boos in patients with sever peripheral edema.

12-Treatment of anemia by EPO.

B-Treatment of Acute hypotension:

1-The patient should be placed in trendlenberg position (if respiratory status allows this ).

2-A bolus of 0.9 saline (100-250 ml) should be administered through the venous line.

Alternative to 0,9 saline :

3% saline :50-100 ml

Dextrose  50 % : 25-50 ml.

Human Albumin 20 % :50-100 ml.

Mannitol.

Pressor agents e,g Dopamine.

3-Rreduce the UF rate near to zero.

4-Discontinue dialysis in severe cases.

.

## Membrane Reactions

(Dialyzer reactions-Membrane Hypersensitivity-First use Syndrome)

Reactions occur with new dialyzer were grouped under the term first use syndrome. However , similar reactions occur with reused dialyzer.These reactions appear to be two varieties :

1-Anaphylactic type (type A)

2-Non specific type (type B)

1-Anaphylactic type (type A)

This type of reaction occur at the first 20-30 minutes of dialysis, usually first 5 minutes. It occur in less than 15% of patients when new dialyzer is used or with 5/100,000 dialyzer sold. Atopic patient are prone to develop these reaction.

Manifestations

Severe reaction: Dyspnea -Sense of impending doom-Heat sensation at fistula site and throughout the body-Cardiac arrest and death may supervene.

Mild reaction: Itching-urticaria - Cough- sneezing-Coryza or watery eye-Abdominal pain and Diarrhea

Etiology

1-Ethylene oxide (used to sterilize HF dialyzers) hypersensitivity, often after inadequate rinising. High IgE antibodies were found in those patient. It is uncommon nowadays.

2-Polyacrylonitrile membrane associated reactions mediated by bradykinin system .

3-ACE inhibitors magnify the effect of the membrane because ACE participate in bradykinin inactivation.

4-Contaminated dialysate with bacteria when high flux dialyzer is used with less frequent cleaning and sterilization.

5-Reuse probably due to bacterial or endotoxin contamination of the water used during the reuse procedure further more formaldehyde and glutaradehyde can cause allergic reaction.

6-Complement activation ; it is conceivable that complement activating membranes and acetate dialysate  may act as co factor in precipitating reaction in some patients, especially those with history of atopy.

7-Heparin allergic reaction should be suspected when other causes reasonably excluded.

6-Others -Azide substances used to package some water pretreatment ultrafiltrers .Acetate causes adenosine release when metabolized can exacerbate bronchoconstriction

Management:

1-Stop dialysis

2-Clamp the blood lines

3-Do not reinfuse the blood

4-Discard the dialyzer and the lines

5-Cardiorespiratory resuscitation

6-Antihistaminic, Steroids ,Epinephrine can be given

Prevention:

1-Proper rinsing of dialyzer prior to use

2-Change suspected ethylene oxide sensitized patient to gamma or steam sterilized dialyzer

3-Stop ACE inhibitor therapy in affected patients.

4-Change the membrane if the reaction occur while using it.

2-Non specific type (type B)

The non specific type of membrane reactions are more common than anaphylactic type but less severe.

Manifestations:

1-It occur within several minutes to one hour

2-Chest pain which may be associated with backache.

Etiology:

Complement activation is the most likely cause .It is less common with reused dialyzer. In reused filters the membrane will be coated with protein layer ,furthermore the toxic substances are washed during the previous dialysis

Prevention

Change the dialyzer membrane to characterized by release of activated complement fragments such as cellulose or Hemophan may be of benefit.

Management:

1-Oxygen therapy

2-Treat as in angina during dialysis

.

## Muscle cramps during hemodialysis

Etiology

1-Hypotension and significant volume of fluid removal by UF

Cramps most commonly occur in association with hypotension & may persist after restoring BP. UF reduce extracellular volume ->>>>>skeletal muscle irritability.

2-Patient below dry weight:

Sever prolonged cramps beginning during the later part of dialysis and persisting after dialysis occur when the patient has been dehydrated to below dry weight.

3-Electrolyte disturbance:

Acute decrease of plasma sodium (due to the use of low sodium dialysate) will result in constriction of blood vesseles, subsequently muscle cramps occur. Hypokalemia also can cause muscle cramps.

Prevention

1- Prevention of hypotension episodes & excessive UF(previous tip).

2-Increasing the dialysis solution sodium level to 140mEq/L or more.

3-Other strategies:\

a) Qunine  sulphate tab 180-300mg ,2 hr. predialysis.

b)Diazepam or Oxazepam   2 hrs pre dialysis .

c)Stretching excercise. to the affected muscle group.

N.B. Side effect of the drugs limit its use.

Management:

1-Concomitant occurrence of hypotension and cramps may respond to treatment with  100 ml  of 0.9 % saline

2-Hypertonic saline or dextrose >>>>>dilatation of the  muscle bed blood vessels and restore blood volume.

*Hypertonic dextrose( 50 ml of  10 % dextrose.) is preferred for treating non diabetic patient to avoid  thirst  produced by saline.

.

## Skin disorders in dialysis patient

a- Related to uremia

pruritus

xerosis

uremic frost

hyperkeratosis penetrans

uremic pigmentation

purpura

calciiphylaxis

bulbous dermatosis

nail changes

b. Drug related

hypertrichosis

acne

hypersensitivity

c. Related to renal disease

cutaneous vasculitis

### Pruritus (Itching)

Pruritis is present at some time in 80-90% of dialyzed patients. It may appear as a symptom of early uremia, but is also troublesome and  persistent in some  patients  on long-term dialysis therapy. Overall, itching was often most severe during or after hemodialysis session but was also increased during periods of inactivity or bed rest.

ETIOLOGY:

1-Circulating uremic toxins.

2-Elevated calcium-phosphorous product.

3-High para thyroid hormone

4-Dry skin aggravate the pruritus.

5-Allergic patients to:

Heparin

Plasticizers used to soften hemodialysis tubing.

Ethylene Oxide gas used  to sterilize hemodialyzers &blood line

DIAGNOSIS:

Pruritis may involve the entire skin surface, predominantly on the face, back, trunk & extremities

Telltale signs of scratching include:

Excoriation , Heemorrhgic crusts , Pustule , Lichenification , Nodule formation

TREATMENT:

A- General measures.

1- Relieve the dry skin with topical emollients e.g. Lubricating solutions, Camphor and Menthol containing creams

2-Switch from ethylene oxide to gamma ray-sterilized dialyzer.

B-Phototherapy:

Ultraviolet light-B delivered in a conventional light box giving 2 treatments per week for 4 weeks..

C-Control of calcium and phosphorous metabolism:

Treat hyperphosphtemia, Control hyperparathyroidism, including vitamin D therapy and partial parathyroidectomy.

D-Drug therapy :   -

Antihistamines,e.g. Diphenhydramine ,Hydroxyzine

Parenteral lidocaine infusion

Oral activated charcoal

E- Renal transplantation

## .

## Causes of Chest pain in ESRD  patients

Cardiovascular

Angina pectoris

Pericarditis

Valvular heart disease Aortic dissection

Pleuropulmonary

Pneumonia

Pleuritis

Hemothorax

Pulmonary embolism/infarction

Gastrointestinal

Esophageal spasm/reflux

Acid peptic disease

Pancreatitis

Cholecystitis/Cholelithiasis

Musculoskeletal

Rib fracture

Renal osteodystrophy

Muscle cramps

Neurologic

Spinal disease with nerve root compression

Herpes zoster

Miscellaneous

Anaphylactic reaction

Air Embolism

## .

## Hemodialysis associated angina

a)Causes:

-Hemodynamic stress caused by hemodialysis

-Reduction of blood Po2 Particularly with acetate dialysate bath and or complement activating cuprophane dialyzers that can cause; pulmonary leukocyte sequestration , change hemoglobin-oxygen affinity, arrhythmia and hypotension (First use syndrome). It occurs in ~5% of dialysis treatment given with fresh unused dialyzers

-Thus hemodialysis reduce coronary artery filling time ,perfusion pressure , and myocardial oxygenation.

-Potential causes of chest pain (e.g. hemolysis ,air embolism) must be considered

-These events coupled with anemia and the possible loss of coronary vasodilator reserve may be responsible for hemodialysis associated angina in patients with or without significant coronary stenosis.

b)Hemodialysis as  a risk factors for angina :

1.Hypotension and Hypertension during dialysis

2.Anemia

3.Hypoxemia during dialysis

4.Chronic volume overload

5.Hyperlipdaemia

6.Acetate intolerance

7.Atherosclerosis

The use of glucose in dialysis fluid may cause hyprlipidemia and thereby atherosclerosis. Acetate dialysis solutions can facilitate synthesis of cholesterol and triglyceride.

c) Management of angina during hemodialysis:

1. Nasal oxygen should be initiated.

2. Reduce blood flow rate and stop ultrafiltration.

3. Sublingual nitroglycerin if blood pressure is maintained.

4. Sedation.

5. Treat the possible cause.

6.Dialysis with bicarbonate dialysate bath.

7.In severe cases discontinue the procedure .

8.ECG and Further investigation may be required to exclude myocardial infarction.

Predialysis administration of nitroglycerin 1houre prior to hemodialysis session may be be of benefit

d) Dialysis after acute myocardial infarction:

1.Postpone dialysis for 24 hours when possible.

2.Such patients are best dialysed peritonealy (if possible) to avoid the attendant     hemodynamic      instability

3.If peritoneal dialysis is impossible -Hemofiltration or Bicarbonate dialysate bath with close monitoring ,oxygen therapy .,blood transfusion if Hct is less than 30% , along with all measures to avoid hypotension (review the previous tips)

.

## Air Embolism

Air embolism is a serious complication of the dialysis procedure. Air can be introduced via the segment  of the blood circuit that operates in the negative pressure range-the portion between the arterial fistula needle and the blood pump.

Routes of air introduction:

a-Air leaks around tubing joints (saline and heparin infusion sites placed before the pump)

b-Excessive undetected or unmonitored negative pressure, related to inadequate fistula flow rate for the pump demand ,or malposition of the arterial needle.

c-Unattended intravenous solution administration.

d-The use of prepump arterial drip chamber

e-The use of air to return blood to the patient at the completion of dialysis.

To avoid that venous  air trap and air detector are located just distal to the venous pressure monitor. The air detector is attached to a relay switch which automatically clamps the venous blood line and shuts off the blood pump if air is detected.

Manifestations

a-In seated patient: Air migrates to cerebral venous system (not to the heart) causing obstruction to cerebral venous return with loss of consciousness , convulsion , even death.

b-In recumbent position: The air tends to enter the heart and generate foam in the right ventricle and pass into the lungs manifested by dyspnea, cough, chest tightness. Further passage of air across the capillary bed into the left ventricle can result in air immobilization to the arteries of the brain and the heart with acute neurological and cardiac dysfunction.

c-Foam will often seen in the venous blood line.

d-A peculiar churning sound may be heard on auscultation ,if the air gone into the heart.

Management

a-Clamp the venous line and stop the blood pump.

b-Place the patient in Trendelenberg position on the left side with the chest and head tilted downward to trap the air at the apex of right ventricle away from the outflow tract.

c-Cardiorespiratory support, Oxygen 100%.

d-Occasionally percutaneous aspiration of the air foam from the heart may be necessary.

e-Other measures include IV Dexamthsone to reduce brain edema ,Heparin /Dextran to improve the microcirculation.

.

## Ruptured  Dialyzer

Causes of dialyzer rupture

1-Improper priming

2-Improper or inadequate heparinization.

3-Damaged dialyzer.

4-Accidental tubing kink; causing high venous pressure

5-Clotting

Management:

Dialyzer rupture is uncommon, if it occurs , an immediate decision must be made to

1- Reinfuse the blood (or not)

2-Change the ruptured dialyzer as quickly as possible.

3-Cleaning the dialysate circuit or reset up.

4-Resume dialysis if safe to do so .

When the rupture is extensive and the pressure in the dialysate compartment is greater than the pressure inside the blood compartment, dialysate can rapidly enter the blood compartment and the patient. To avoid that, the machine should always be set to at least 25-50 mm Hg negative pressure.

[Usually less than 1% of blood present in dialysate trigger the blood leak alarm

.

## Clotted dialyzer

Causes

a. High venous Pressure

b. Slow blood flow

c.Turbulence of air that may have been infused during priming

eTurbulence of air and blood while using single needle machine

f. Blood transfusion during dialysis

Mechanism:

Platelet activation upon contact with the dialyzer membrane is the primary mechanism responsible for clotting during dialysis.

Management:

When a dialyzer clots , the dialyzer, the arterial line, the venous line may need to be replaced. In clotted dialyzer without rupture, the entire dialysate circuit does not need to be set up again or recleaned

Dialyzers need not to clot even if no heparin is used during dialysis.The procedure consisted of clamping the arterial line every 20-30 minutes, and rinsing the blood circuit with normal saline to clear the dialyzer and blood lines. Negative pressure  to be adjusted to compensate  for the excess fluid  administered. Saline rinsing interferes with platelet aggregation and prevents overproduction of active coagulation factors. Thus the formation of stable fibrin network is prevented.

.

## Needle-Site Bleeding

Needle-site bleeding may occur after venipuncture as a result of:

1-capillary trauma or trauma occurring when rotating the needle or manipulating the needle side to side during insertion.

2-Multiple puncture at the same site

3-Overheparinization

Management:

1-Direct pressure is the simplest and effective measure. It should be aseptically and with extreme care to avoid AV access occlusion.

2-Apply gel foam to the puncture site

3-Adjust heparin dose

4-Change the puncture site every dialysis

.

## High Venous Pressure

High venous pressure is seen on almost daily basis. The purpose of monitoring venous pressure is to identify any pressure changes within venous return line during dialysis. Venous pressure of +50 to +100 mmHg above the atmospheric pressure is acceptable during dialysis(preset pressure -100 to +100 ). If venous pressure moves out of limits an audible alarm is heard and the blood pump stopped until the problem is corrected.

Causes of high venous pressure:

1-Needle related –

Needle may have infiltrated the access wall

Needle hole may be resting against the side wall of the access

Needle may have been inserted into scar tissue inside the access

Needle may be to short and the hole is not completely inside the access

Needle with small gouge (16) when blood flow rate is high.

2-Clotting

Clotting of venous needle

Clotting of venous side of AV shunt

Clotting of venous blood line

Clotting of the filter

3-Access related

Stenosis or spasm at the venous limb of the access

Clotting at the venous limb of the access

AV graft have naturally high venous pressure(150-200mmHg)

4-Lines related

Kinked venous line

Management  ( See also alarm problem solving)

1-Manipulate the needle and or the line. If the access is small, a tourniquet must be used, being certain the blood pump is off.. Recannulation with new needle if needed ,the original one should be left in place until the end of dialysis. to avoid undue bleeding as the patient is heparinized

2-Adjust the blood flow rate, Proper heparnization, Treat access problem, and Proper needle and needling

3-Extreme care must be exercised when dialyzing a patient with high venous pressure

This increases the baseline TMP and obligatory ultrafiltration will occur

Single-needle device is occasionally impossible to use, because with high venouspressure ,venous return will be impaired and blood recirculation will be high.

.

## Abnormal Arterial Pressure

The arterial pressure monitor reads in negative pressure because it measures the partial vacuum resulting from the roller pump withdrawal of the whole blood at a relatively high rate through the small bore needle inserted into the arterial end of the access. It identifies how much suction is being placed on the arterial wall. When the pressure becomes very negative (-150 &-200mmHg) the access will collapse and the blood flow will be discontinued to the machine. Furthermore very high negative pressure may cause air bubbles to enter around the needle causing air embolism and or clotting of the dialyzer.

Causes of an increase  in negative pressure:

1-Needle related –

Needle may have infiltrated the access wall

Needle hole may be resting against the side wall of the access

Needle may have been inserted into scar tissue or flap inside the access

2-Clotting

Clotted arterial needle

Clotted arterial side of AV shunt

3-Access related  -

Stenosis within the access

Spasm of the access.

Positional in acute access

4-Patient related   -Low blood pressure

Management:   ( See also alarm problem solving)

1 Manipulation of the arterial needle is similar to that of venous needle. In most cases the needle may have to be replaced with the same precaution as with venous needle

2-Treat the cause.

3- Proper needling

4-Asses the access

.

## Hemolysis

Acute hemolysis during hemodialysis is rare  but must be immediately identified and emergency management initiated to minimize the patient morbidity.

Technical causes:

1-Hypotonic dialysate

2-Overheated dialysate

3-Formaldehyde in the dialysate

4-Hypochlorite in dialysate

5-Copper or Nitrates in the dialysate

6-Overoccluded blood pump

7-Severe foaming in the roller clamp segment

Manifestations

1-Port-wine appearance of blood in the venous return line

2-Presence of pink plasma in centrifuged blood sample.

3-Patient Manifestations :Headache , Arm pain , Chest pain, Drowsiness Dyspnea , Back pain, Diaphoresis  ,  Pallor, Arrhythmia ,Cardiac arrest

4-Laboratory studies, Drop in Hemoglobin and Hematocrit ,Hyperkalemia, Arrhythmias ,Cardiac arrest

Management

1-Clamp the blood lines

2-Turn the blood pump off

3-Do not reunifies  the blood

4-Leave the needle in place

5-Draw blood for: Type and crossmatch    CBC    Electrolyte  Trace metal

6-Save dialysate and blood samples for further evaluation

7-Discard the dialyzer lines

8-Evaluate the problem by reviewing the following

a. Check conductivity

b.Check the dialysate temperature

c. Look for kinks in the blood line in the pump

d. Check dialysate for ;Electrolyte Formaldehyde-Chlorine-Trace metal

9-Check the IV priming solution

!0-Resume dialysis to prevent hyperkalemia and blood transfusion  if necessary

.

## Conductivity

Conductivity is a property of electrolyte solutions to conduct an electric current. It is measured in units called mhos or milli mohos (1mhos =1000 milli mhos). A mho is the conductivity of body through which 1 A of current flows when potential difference is 1 V. It is important to be aware that the display meter of clinical dialysis machines does not read in milli mhos, but displays an entirely arbitrary scale intended to present a clinically understandable reading approximately equal to mEq/L chloride ions  in the dialysate solution. Because dialysate chloride  concentration is usually 105 mEq/L, the scale deviation from the ideal set line are, therefore, approximately % deviation from the desired reading. All recent dialysis machine have factory-set conductivity meters with alarm system. When conductivity meter detects high or low conductivity  the machine automatically sounds an alarm and puts the dialysate into bypass mode so that  no dialysate flow to the dialyzer.

Usually conductivity is adjusted between 13-14 in single bath system. In Redy Sorbent hemodialysis the initial dialysis solution conductivity adjusted according to patient natremic status (it will be discussed in later tips). The newer computerized dialysis machine display the actual sodium and electrolyte concentration in dialysate (profiling)

High-Conductivity Alarm:

Causes : High-conductivity is a dialysate production complication caused by too much concentrate and not enough treated water

Management :   (See also alarm problem solving)

Evaluate the incoming water and check for kink the line.

The water pressure and filters should be also checked

Recheck the conductivity:

If normal, there is malfunction in the machine itself (change it)

If high again concentrate should be changed

Check the conductivity before resuming dialysis

Low-Conductivity Alarm:

Causes : Low conductivity is a dialysate machine mixing complication that occurs when the machine receives insufficient dialysate concentrate within the mixing chamber. The result is a hypotonic dialysate that cause hemolysis. The usual reasons are :

1-The concentrate container runs dry

2-Leakage at the concentrate pump .

3-Malfunction in the concentrate pump feedback circuit

Management : (See also alarm problem solving)

Change the concentrate container if it is dry

If the concentrate container has not run dry, a sample of dialysate should sent to the      laboratory for sodium and chloride

Recheck the conductivity with second conductivity meter:

If low conductivity is confirmed  ; change the concentrate bottle (recheck again)

If the conductivity still low after changing the concentrate , a different machine should be  tried

# Comments

Reference www.dialysistips.com

# Text boxes

Native AV fistula

Loop AV graft